## GROUP 1 CC STANDARD
DATA FOR DATACOLOR SPECTRAFLASH 450 | [12 empty] | DATA FOR DATACOLOR SPECTRAFLASH 600 | [12 empty] | DATA FOR DATACOLOR SPECTRAFLASH 300
Inst, no: | 435 | 435 | 435 | 435 | 435 | 435 | 435 | 435 | 435 | 435 | 435 | 435 | Inst, no: | 596 | 596 | 596 | 596 | 596 | 596 | 596 | 596 | 596 | 596 | 596 | 596 | Inst, no: | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118
Conditions: | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | Conditions: | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | Conditions: | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib
Name: | medium blue | dark blue | greenish blue | dark green | light green | yellow | brown | pink | maroon | dark grey | medium grey | light grey | Name: | medium blue | dark blue | greenish blue | dark green | light green | yellow | brown | pink | maroon | dark grey | medium grey | light grey | Name: | medium blue | dark blue | greenish blue | dark green | light green | yellow | brown | pink | maroon | dark grey | medium grey | light grey
X | 15.41 | 6.12 | 9.59 | 8.74 | 29.06 | 59.17 | 23.39 | 48.18 | 7.19 | 9.02 | 27.69 | 59.08 | X | 15.23 | 5.88 | 9.35 | 8.51 | 28.88 | 59.28 | 23.23 | 48.22 | 6.95 | 8.77 | 27.55 | 59.27 | X | 15.28 | 6.01 | 9.44 | 8.6 | 28.89 | 59.02 | 23.04 | 47.85 | 6.99 | 8.86 | 27.5 | 58.92
Y | 16.75 | 5.73 | 11.82 | 10.63 | 34.65 | 62.99 | 19.57 | 44.95 | 6.26 | 9.54 | 29.79 | 62.99 | Y | 16.56 | 5.47 | 11.58 | 10.39 | 34.47 | 63.03 | 19.37 | 44.98 | 6 | 9.27 | 29.64 | 63.19 | Y | 16.61 | 5.62 | 11.65 | 10.46 | 34.44 | 62.69 | 19.28 | 44.72 | 6.08 | 9.37 | 29.58 | 62.81
Z | 27.83 | 11.5 | 14.47 | 7.53 | 28.39 | 16.05 | 8.71 | 41.21 | 6.01 | 10.86 | 32.64 | 64.56 | Z | 27.64 | 11.19 | 14.08 | 7.17 | 27.97 | 15.73 | 8.32 | 41.03 | 5.62 | 10.47 | 32.37 | 64.68 | Z | 27.67 | 11.38 | 14.21 | 7.38 | 27.98 | 15.84 | 8.54 | 40.79 | 5.83 | 10.66 | 32.36 | 64.31
x | 0.2569 | 0.262 | 0.2673 | 0.3248 | 0.3155 | 0.4281 | 0.4526 | 0.3586 | 0.3695 | 0.3066 | 0.3073 | 0.3166 | x | 0.2563 | 0.261 | 0.2672 | 0.3263 | 0.3163 | 0.4294 | 0.4562 | 0.3593 | 0.374 | 0.3075 | 0.3077 | 0.3167 | x | 0.2565 | 0.2612 | 0.2674 | 0.3252 | 0.3164 | 0.4291 | 0.4531 | 0.3588 | 0.3697 | 0.3067 | 0.3075 | 0.3167
y | 0.2792 | 0.2453 | 0.3295 | 0.3952 | 0.3762 | 0.4558 | 0.3788 | 0.3346 | 0.3218 | 0.3242 | 0.3306 | 0.3375 | y | 0.2786 | 0.2425 | 0.3307 | 0.3987 | 0.3775 | 0.4566 | 0.3804 | 0.3351 | 0.3231 | 0.3252 | 0.3309 | 0.3377 | y | 0.2789 | 0.2443 | 0.3301 | 0.3957 | 0.3772 | 0.4558 | 0.379 | 0.3353 | 0.3218 | 0.3244 | 0.3307 | 0.3376
Wcie | 150.32 | 192.77 | 51.66 | -107.2 | -43.56 | -240.5 | -172.7 | 2.98 | -22.56 | 26.9 | 35.77 | 49.73 | Wcie | 151.65 | 198.1 | 49.36 | -114.7 | -46.46 | -242.9 | -178.5 | 1.67 | -28.78 | 24.23 | 34.67 | 49.54 | Wcie | 151.05 | 195.16 | 50.41 | -108.7 | -46.13 | -241.6 | -173.7 | 1.38 | -22.91 | 26.32 | 35.17 | 49.23
Tcie | 17.56 | -9.08 | 40.88 | 31.82 | 27.85 | -21.76 | -93.8 | -37.98 | -56.11 | 2.08 | 5.61 | 1.77 | Tcie | 17.71 | -9.98 | 41.83 | 32.75 | 28.02 | -22.42 | -96.02 | -38.22 | -59.29 | 1.86 | 5.51 | 1.74 | Tcie | 17.7 | -9 | 41.19 | 31.82 | 27.68 | -22.64 | -94.12 | -37.68 | -56.25 | 2.13 | 5.48 | 1.7
Wgg | 320.11 | 424.94 | 109.62 | -242.6 | -130.9 | -608.1 | -412.4 | -47.31 | -58.77 | 53.28 | 48.64 | 38.73 | Wgg | 322.54 | 436.24 | 104.09 | -259.8 | -138.1 | -614.6 | -426.3 | -51.18 | -73.47 | 46.82 | 45.53 | 37.18 | Wgg
Tgg | 22.35 | -7.25 | 47.87 | 36.31 | 32.17 | -25.73 | -106 | -41.82 | -62.16 | 3.9 | 7.78 | 3.29 | Tgg | 22.43 | -8.3 | 48.75 | 37 | 32.07 | -26.94 | -109 | -42.36 | -66.06 | 3.42 | 7.46 | 3.02 | Tgg
L* | 47.94 | 28.71 | 40.93 | 38.95 | 65.47 | 83.44 | 51.35 | 72.86 | 30.07 | 37 | 61.48 | 83.44 | L* | 47.69 | 28.03 | 40.54 | 38.54 | 65.34 | 83.46 | 51.12 | 72.88 | 29.41 | 36.49 | 61.34 | 83.54 | L* | 47.76 | 28.44 | 40.66 | 38.66 | 65.31 | 83.28 | 51.01 | 72.7 | 29.62 | 36.69 | 61.29 | 83.34
a* | -2.74 | 7.81 | -12.42 | -11.01 | -14.06 | -1.33 | 23.26 | 15.98 | 13.1 | -0.19 | -2.19 | -1.54 | a* | -2.75 | 8.19 | -12.67 | -11.24 | -14.16 | -1.14 | 23.57 | 16.02 | 13.46 | -0.18 | -2.18 | -1.54 | a* | -2.76 | 7.83 | -12.48 | -10.96 | -14.01 | -1 | 23.2 | 15.74 | 13 | -0.21 | -2.15 | -1.51
b* | -17.29 | -17.91 | -4.4 | 12.24 | 12.08 | 65.29 | 29.53 | 7.83 | 2.91 | -1.83 | -0.93 | 2.6 | b* | -17.43 | -18.23 | -4.14 | 12.89 | 12.47 | 66.01 | 30.45 | 8.08 | 3.45 | -1.56 | -0.78 | 2.68 | b* | -17.36 | -18.06 | -4.26 | 12.31 | 12.42 | 65.47 | 29.51 | 8.06 | 2.89 | -1.78 | -0.87 | 2.66
Name: | medium blue | dark blue | greenish blue | dark green | light green | yellow | brown | pink | maroon | dark grey | medium grey | light grey | Name: | medium blue | dark blue | greenish blue | dark green | light green | yellow | brown | pink | maroon | dark grey | medium grey | light grey | Name: | medium blue | dark blue | greenish blue | dark green | light green | yellow | brown | pink | maroon | dark grey | medium grey | light grey
360 | 16.18 | 11.99 | 8.8 | 4.98 | 10.06 | 6.31 | 6.45 | 28.79 | 6.67 | 8.67 | 20.75 | 40.76 | 360 | [12 empty] | 360
370 | 19.13 | 13.66 | 10.21 | 5.32 | 11.03 | 6.74 | 7.05 | 29.69 | 7.08 | 9.43 | 23.3 | 46.02 | 370 | [12 empty] | 370
380 | 22.12 | 15.32 | 11.27 | 5.59 | 11.87 | 6.85 | 7.18 | 27.89 | 6.94 | 9.81 | 25.21 | 50.24 | 380 | [12 empty] | 380
390 | 24.66 | 16.81 | 11.98 | 6.64 | 12.91 | 7.07 | 7.17 | 25.99 | 6.78 | 10.06 | 26.76 | 53.38 | 390 | [12 empty] | 390
400 | 26.62 | 17.75 | 12.47 | 9.28 | 14.29 | 7.46 | 7.27 | 25.39 | 6.75 | 10.29 | 27.97 | 55.49 | 400 | 26.49 | 17.45 | 12.02 | 8.78 | 13.97 | 7.29 | 6.84 | 25.43 | 6.29 | 9.87 | 27.84 | 55.88 | 400 | 26.43 | 17.38 | 12.16 | 8.84 | 14.13 | 7.36 | 7.07 | 25.37 | 6.45 | 10.07 | 27.79 | 55.47
410 | 27.78 | 17.6 | 12.65 | 10.48 | 16.32 | 7.89 | 7.37 | 26.38 | 6.54 | 10.34 | 28.97 | 57.11 | 410 | 27.58 | 17.31 | 12.19 | 10.15 | 15.86 | 7.66 | 6.94 | 26.22 | 6.1 | 9.95 | 28.68 | 57.25 | 410 | 27.55 | 17.28 | 12.36 | 10.16 | 16.04 | 7.79 | 7.18 | 26.18 | 6.29 | 10.16 | 28.7 | 56.94
420 | 28.15 | 16.42 | 12.69 | 9.19 | 18.53 | 8.48 | 7.59 | 27.44 | 6.25 | 10.36 | 29.59 | 58.2 | 420 | 27.96 | 16.18 | 12.27 | 8.86 | 18.09 | 8.26 | 7.17 | 27.31 | 5.85 | 9.96 | 29.29 | 58.27 | 420 | 27.94 | 16.21 | 12.45 | 9.05 | 18.2 | 8.37 | 7.39 | 27.25 | 6.04 | 10.17 | 29.33 | 57.95
430 | 28.22 | 14.68 | 12.93 | 7.71 | 20.71 | 9.3 | 7.85 | 28.77 | 6.01 | 10.37 | 30.08 | 59.05 | 430 | 28 | 14.36 | 12.51 | 7.34 | 20.27 | 9.03 | 7.46 | 28.59 | 5.59 | 9.95 | 29.79 | 59.08 | 430 | 28.01 | 14.48 | 12.67 | 7.58 | 20.35 | 9.16 | 7.66 | 28.56 | 5.81 | 10.14 | 29.79 | 58.74
440 | 27.76 | 12.55 | 13.15 | 6.75 | 22.94 | 10.37 | 8.02 | 31.92 | 5.74 | 10.28 | 30.34 | 59.56 | 440 | 27.61 | 12.28 | 12.77 | 6.41 | 22.52 | 10.1 | 7.66 | 31.57 | 5.38 | 9.91 | 30.08 | 59.61 | 440 | 27.61 | 12.45 | 12.93 | 6.63 | 22.6 | 10.26 | 7.89 | 31.56 | 5.61 | 10.11 | 30.1 | 59.28
450 | 26.89 | 10.54 | 13.4 | 6.32 | 25.5 | 11.97 | 8.14 | 37.63 | 5.56 | 10.21 | 30.56 | 60.06 | 450 | 26.72 | 10.25 | 13.03 | 5.98 | 25.13 | 11.71 | 7.79 | 37.4 | 5.21 | 9.86 | 30.31 | 60.21 | 450 | 26.75 | 10.45 | 13.15 | 6.18 | 25.12 | 11.84 | 7.98 | 37.1 | 5.41 | 10.03 | 30.3 | 59.85
460 | 25.59 | 8.77 | 13.66 | 6.18 | 28.46 | 14.46 | 8.09 | 43.71 | 5.41 | 10.09 | 30.68 | 60.66 | 460 | 25.43 | 8.49 | 13.32 | 5.84 | 28.06 | 14.14 | 7.72 | 43.51 | 5.06 | 9.73 | 30.44 | 60.78 | 460 | 25.49 | 8.73 | 13.42 | 6.05 | 28.03 | 14.26 | 7.93 | 43.12 | 5.25 | 9.9 | 30.43 | 60.45
470 | 23.99 | 7.52 | 13.88 | 6.25 | 31.41 | 17.98 | 7.99 | 47.9 | 5.29 | 9.91 | 30.76 | 61.1 | 470 | 23.8 | 7.2 | 13.56 | 5.91 | 31.05 | 17.65 | 7.62 | 47.81 | 4.94 | 9.56 | 30.5 | 61.25 | 470 | 23.85 | 7.44 | 13.65 | 6.1 | 31 | 17.8 | 7.85 | 47.45 | 5.12 | 9.72 | 30.49 | 60.88
480 | 22.32 | 6.68 | 14.18 | 6.58 | 34.39 | 23.66 | 8.21 | 50.35 | 5.2 | 9.74 | 30.8 | 61.68 | 480 | 22.16 | 6.39 | 13.83 | 6.22 | 33.93 | 23.04 | 7.84 | 50.35 | 4.86 | 9.4 | 30.54 | 61.79 | 480 | 22.24 | 6.61 | 13.92 | 6.39 | 33.86 | 23.11 | 8.04 | 49.96 | 5.03 | 9.56 | 30.52 | 61.44
490 | 21.1 | 6.23 | 14.44 | 7.17 | 36.52 | 30.53 | 8.7 | 49.87 | 5.11 | 9.6 | 30.76 | 62.28 | 490 | 20.87 | 5.91 | 14.15 | 6.88 | 36.33 | 30.48 | 8.39 | 49.95 | 4.79 | 9.28 | 30.55 | 62.46 | 490 | 20.94 | 6.13 | 14.22 | 7.01 | 36.18 | 30.27 | 8.55 | 49.63 | 4.95 | 9.43 | 30.5 | 62.04
500 | 20.08 | 5.94 | 14.85 | 8.08 | 38.35 | 40.01 | 9.59 | 46.86 | 5.05 | 9.53 | 30.74 | 62.91 | 500 | 19.88 | 5.65 | 14.57 | 7.8 | 38.13 | 39.82 | 9.3 | 46.93 | 4.76 | 9.23 | 30.56 | 63.1 | 500 | 19.96 | 5.85 | 14.63 | 7.94 | 37.99 | 39.49 | 9.43 | 46.74 | 4.91 | 9.37 | 30.51 | 62.69
510 | 19.29 | 5.75 | 15.38 | 9.58 | 39.42 | 49.97 | 10.8 | 42.47 | 5.03 | 9.52 | 30.72 | 63.39 | 510 | 19.09 | 5.47 | 15.07 | 9.31 | 39.19 | 49.6 | 10.51 | 42.58 | 4.75 | 9.22 | 30.53 | 63.57 | 510 | 19.16 | 5.66 | 15.1 | 9.42 | 39.05 | 49.17 | 10.61 | 42.5 | 4.89 | 9.35 | 30.47 | 63.17
520 | 18.87 | 5.67 | 15.57 | 12.56 | 39.63 | 58.09 | 12.27 | 38.48 | 5.04 | 9.6 | 30.64 | 63.68 | 520 | 18.65 | 5.37 | 15.31 | 12.24 | 39.4 | 57.89 | 11.97 | 38.48 | 4.75 | 9.3 | 30.43 | 63.83 | 520 | 18.68 | 5.55 | 15.29 | 12.23 | 39.29 | 57.42 | 12.03 | 38.44 | 4.88 | 9.41 | 30.36 | 63.44
530 | 18.62 | 5.64 | 15.12 | 15.43 | 39.1 | 63.92 | 13.83 | 35.47 | 5.06 | 9.71 | 30.46 | 63.77 | 530 | 18.42 | 5.37 | 14.89 | 15.24 | 38.94 | 63.89 | 13.59 | 35.49 | 4.79 | 9.44 | 30.3 | 63.97 | 530 | 18.43 | 5.53 | 14.9 | 15.08 | 38.87 | 63.47 | 13.6 | 35.44 | 4.91 | 9.54 | 30.23 | 63.58
540 | 18.26 | 5.66 | 13.99 | 15.82 | 38.03 | 67.71 | 15.37 | 33.89 | 5.13 | 9.83 | 30.23 | 63.67 | 540 | 18.09 | 5.4 | 13.77 | 15.67 | 37.9 | 67.84 | 15.19 | 33.88 | 4.88 | 9.58 | 30.11 | 63.88 | 540 | 18.1 | 5.56 | 13.83 | 15.59 | 37.87 | 67.45 | 15.17 | 33.84 | 4.99 | 9.67 | 30.05 | 63.52
550 | 17.43 | 5.65 | 12.48 | 14.03 | 36.73 | 70.36 | 16.93 | 33.81 | 5.28 | 9.87 | 30.05 | 63.53 | 550 | 17.24 | 5.38 | 12.24 | 13.79 | 36.57 | 70.46 | 16.76 | 33.78 | 5.02 | 9.61 | 29.92 | 63.77 | 550 | 17.27 | 5.53 | 12.35 | 13.88 | 36.56 | 70.08 | 16.7 | 33.7 | 5.13 | 9.7 | 29.85 | 63.37
560 | 15.87 | 5.48 | 10.82 | 11.53 | 35.22 | 72.19 | 18.61 | 35.3 | 5.47 | 9.67 | 29.85 | 63.33 | 560 | 15.71 | 5.23 | 10.62 | 11.34 | 35.09 | 72.25 | 18.4 | 35.23 | 5.21 | 9.42 | 29.71 | 63.55 | 560 | 15.79 | 5.38 | 10.74 | 11.49 | 35.11 | 71.93 | 18.33 | 35.12 | 5.32 | 9.52 | 29.64 | 63.16
570 | 14.29 | 5.28 | 9.6 | 9.62 | 33.77 | 73.28 | 20.54 | 38.58 | 5.74 | 9.38 | 29.64 | 63.16 | 570 | 14.09 | 5.03 | 9.38 | 9.41 | 33.62 | 73.39 | 20.33 | 38.5 | 5.49 | 9.12 | 29.47 | 63.32 | 570 | 14.2 | 5.18 | 9.5 | 9.56 | 33.63 | 73.05 | 20.21 | 38.28 | 5.58 | 9.22 | 29.4 | 62.94
580 | 13.18 | 5.14 | 8.92 | 8.46 | 32.4 | 73.76 | 23.01 | 43.82 | 6.13 | 9.14 | 29.35 | 62.91 | 580 | 12.97 | 4.9 | 8.69 | 8.23 | 32.26 | 73.96 | 22.86 | 43.8 | 5.9 | 8.9 | 29.21 | 63.13 | 580 | 13.07 | 5.05 | 8.79 | 8.36 | 32.29 | 73.58 | 22.68 | 43.44 | 5.96 | 8.99 | 29.13 | 62.75
590 | 12.61 | 5.09 | 8.44 | 7.91 | 31.24 | 74.43 | 26.38 | 51.27 | 6.71 | 9.05 | 29.07 | 62.8 | 590 | 12.41 | 4.86 | 8.22 | 7.67 | 31.1 | 74.66 | 26.23 | 51.28 | 6.47 | 8.81 | 28.95 | 63.03 | 590 | 12.47 | 4.99 | 8.29 | 7.78 | 31.11 | 74.27 | 25.99 | 50.81 | 6.53 | 8.89 | 28.87 | 62.65
600 | 12.39 | 5.11 | 8.04 | 7.78 | 30.26 | 75.52 | 30.76 | 60.39 | 7.53 | 9.06 | 28.77 | 62.84 | 600 | 12.2 | 4.88 | 7.81 | 7.55 | 30.12 | 75.83 | 30.64 | 60.53 | 7.3 | 8.83 | 28.65 | 63.03 | 600 | 12.24 | 5 | 7.9 | 7.64 | 30.15 | 75.45 | 30.3 | 59.85 | 7.33 | 8.91 | 28.6 | 62.66
610 | 12.29 | 5.14 | 7.8 | 7.91 | 29.42 | 76.59 | 36.15 | 69.07 | 8.66 | 9.11 | 28.42 | 62.71 | 610 | 12.1 | 4.9 | 7.57 | 7.69 | 29.32 | 76.88 | 36.03 | 69.16 | 8.44 | 8.88 | 28.34 | 62.94 | 610 | 12.15 | 5.02 | 7.66 | 7.75 | 29.35 | 76.5 | 35.48 | 68.31 | 8.41 | 8.95 | 28.27 | 62.58
620 | 12.3 | 5.16 | 7.82 | 8.27 | 28.85 | 77.18 | 41.53 | 74.52 | 10.18 | 9.24 | 28.12 | 62.56 | 620 | 12.08 | 4.92 | 7.58 | 8.03 | 28.71 | 77.51 | 41.39 | 74.78 | 9.91 | 8.99 | 28 | 62.82 | 620 | 12.11 | 5.02 | 7.63 | 8.06 | 28.73 | 77.12 | 40.72 | 74.04 | 9.84 | 9.04 | 27.95 | 62.46
630 | 12.26 | 5.15 | 7.83 | 8.74 | 28.44 | 77.5 | 45.15 | 76.95 | 11.98 | 9.33 | 27.78 | 62.47 | 630 | 12.05 | 4.94 | 7.61 | 8.51 | 28.29 | 77.88 | 45.15 | 77.41 | 11.73 | 9.1 | 27.66 | 62.7 | 630 | 12.09 | 5.03 | 7.66 | 8.52 | 28.31 | 77.55 | 44.46 | 76.74 | 11.6 | 9.14 | 27.6 | 62.33
640 | 12.23 | 5.17 | 7.79 | 9.39 | 28.2 | 77.99 | 45.59 | 76.73 | 14.1 | 9.4 | 27.41 | 62.27 | 640 | 12.03 | 4.96 | 7.57 | 9.17 | 28.06 | 78.27 | 45.96 | 77.29 | 13.88 | 9.18 | 27.31 | 62.5 | 640 | 12.09 | 5.06 | 7.64 | 9.19 | 28.08 | 77.95 | 45.29 | 76.77 | 13.7 | 9.24 | 27.26 | 62.16
650 | 12.34 | 5.22 | 7.76 | 10.32 | 28.11 | 78.58 | 42.4 | 74.8 | 16.42 | 9.5 | 27.02 | 61.9 | 650 | 12.14 | 5.01 | 7.54 | 10.09 | 27.99 | 78.91 | 42.72 | 75.29 | 16.22 | 9.28 | 26.95 | 62.21 | 650 | 12.19 | 5.11 | 7.61 | 10.1 | 28.02 | 78.59 | 42.47 | 75 | 15.99 | 9.34 | 26.91 | 61.89
660 | 12.78 | 5.33 | 7.89 | 11.51 | 28.22 | 79.37 | 39.4 | 74.17 | 18.8 | 9.72 | 26.63 | 61.77 | 660 | 12.59 | 5.14 | 7.69 | 11.31 | 28.11 | 79.71 | 39.58 | 74.48 | 18.68 | 9.52 | 26.57 | 62.02 | 660 | 12.6 | 5.22 | 7.74 | 11.27 | 28.12 | 79.41 | 39.38 | 74.25 | 18.36 | 9.56 | 26.54 | 61.74
670 | 13.81 | 5.59 | 8.47 | 12.86 | 28.5 | 80.27 | 38.47 | 75.15 | 21.15 | 10.22 | 26.22 | 61.76 | 670 | 13.59 | 5.41 | 8.24 | 12.67 | 28.37 | 80.58 | 38.63 | 75.58 | 21.07 | 10.01 | 26.16 | 61.98 | 670 | 13.57 | 5.47 | 8.26 | 12.6 | 28.35 | 80.2 | 38.27 | 75.09 | 20.69 | 10.02 | 26.11 | 61.66
680 | 15.82 | 6.23 | 9.93 | 14.29 | 28.95 | 81.16 | 37.33 | 76.18 | 23.45 | 11.18 | 25.8 | 61.71 | 680 | 15.54 | 6.04 | 9.65 | 14.07 | 28.79 | 81.44 | 37.56 | 76.46 | 23.34 | 10.94 | 25.73 | 61.82 | 680 | 15.45 | 6.07 | 9.6 | 13.98 | 28.77 | 81.11 | 37.31 | 76.08 | 22.9 | 10.92 | 25.71 | 61.56
690 | 18.49 | 7.69 | 12.23 | 15.41 | 29.18 | 81.65 | 35.15 | 77.78 | 24.94 | 12.29 | 25.18 | 61.13 | 690 | 18.53 | 7.54 | 12.15 | 15.48 | 29.39 | 82.32 | 35.51 | 78.46 | 25.3 | 12.26 | 25.37 | 61.67 | 690 | 18.41 | 7.56 | 12.06 | 15.4 | 29.35 | 81.87 | 35.42 | 77.93 | 24.85 | 12.21 | 25.33 | 61.37
700 | 21.81 | 10.45 | 15.18 | 16.74 | 29.79 | 82.18 | 35.26 | 80.33 | 26.41 | 13.67 | 24.88 | 61.08 | 700 | 22.54 | 11.1 | 15.7 | 17.07 | 30.11 | 82.97 | 36.06 | 81.39 | 27.04 | 13.91 | 24.98 | 61.55 | 700 | 22.32 | 10.9 | 15.48 | 16.97 | 30.05 | 82.55 | 35.76 | 80.78 | 26.57 | 13.82 | 24.96 | 61.26

## GROUP 2 OO STANDARD
DATA FOR DATACOLOR SPECTRAFLASH 450 | [13 empty] | DATA FOR DATACOLOR SPECTRAFLASH 600 | [12 empty] | DATA FOR DATACOLOR SPECTRAFLASH 300
Inst, no: | 435 | 435 | 435 | 435 | 435 | 435 | 435 | 435 | 435 | 435 | 435 | 435 | 435 | 435 | 435 | 435 | Inst, no: | 596 | 596 | 596 | 596 | 596 | 596 | 596 | 596 | 596 | 596 | 596 | 596 | 596 | 596 | 596 | 596 | Inst, no: | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118
Conditions: | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | Conditions: | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | Conditions: | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib
Name: | 00 white | 01 gray | 02 bright red | 03 lidka red | 04 dark pink | 05 gosia pink | 06 yellow | 07 adam olive | 07 yellow olive | 09 yellow beige | 10 green | 11 forest green (dark) | 12 mint | 13 dark blue | 14 blue | 15 light blue | Name: | 00 white | 01 gray | 02 bright red | 03 lidka red | 04 dark pink | 05 gosia pink | 06 yellow | 07 adam olive | 08 yellow olive | 09 yellow beige | 10 green | 11 forest green (dark) | 12 mint | 13 dark blue | 14 blue | 15 light blue | Name: | 00 white | 01 gray | 02 bright red | 03 lidka red | 04 dark pink | 05 gosia pink | 06 yellow | 07 adam olive | 08 yellow olive | 09 yellow beige | 10 green | 11 forest green (dark) | 12 mint | 13 dark blue | 14 blue | 15 light blue
X | 79.05 | 10.37 | 12.69 | 8.17 | 31.7 | 59.01 | 54.23 | 11.82 | 26.97 | 67.16 | 8.31 | 7.14 | 52.48 | 8.28 | 20.03 | 39.24 | X | 79.28 | 10.23 | 12.49 | 7.92 | 31.6 | 58.97 | 54.34 | 11.65 | 26.87 | 67.04 | 8 | 6.85 | 52.39 | 8.08 | 19.84 | 39.17 | X | 79.17 | 10.24 | 12.35 | 8 | 31.31 | 58.79 | 54.15 | 11.73 | 26.83 | 66.86 | 8.17 | 7.02 | 52.25 | 8.21 | 19.93 | 39.14
Y | 84.11 | 10.91 | 8.74 | 6.6 | 27.5 | 56.81 | 51.91 | 12.26 | 27.83 | 71.33 | 11.86 | 8.57 | 60.86 | 10.59 | 24.66 | 45.39 | Y | 84.34 | 10.76 | 8.52 | 6.36 | 27.35 | 56.71 | 51.99 | 12.07 | 27.69 | 71.16 | 11.55 | 8.27 | 60.73 | 10.4 | 24.48 | 45.33 | Y | 84.21 | 10.77 | 8.51 | 6.47 | 27.12 | 56.57 | 51.77 | 12.15 | 27.62 | 70.94 | 11.69 | 8.43 | 60.48 | 10.55 | 24.59 | 45.3
Z | 91.21 | 12.18 | 5.29 | 5.45 | 21.47 | 58.86 | 5.8 | 5.7 | 7.39 | 47.47 | 9.33 | 8.62 | 55.26 | 22.82 | 42.28 | 64.21 | Z | 91.22 | 11.95 | 4.93 | 5.07 | 21.17 | 58.63 | 5.54 | 5.38 | 7.09 | 47.06 | 8.84 | 8.16 | 54.93 | 22.52 | 42.06 | 64.15 | Z | 91.26 | 12.05 | 5.14 | 5.28 | 21.04 | 58.62 | 5.71 | 5.57 | 7.36 | 47 | 9.07 | 8.45 | 54.72 | 22.61 | 42.04 | 63.97
x | 0.3108 | 0.3099 | 0.475 | 0.4039 | 0.393 | 0.3378 | 0.4844 | 0.3969 | 0.4337 | 0.3611 | 0.2816 | 0.2934 | 0.3113 | 0.1987 | 0.2303 | 0.2636 | x | 0.3111 | 0.3107 | 0.4816 | 0.4093 | 0.3944 | 0.3383 | 0.4858 | 0.4004 | 0.4358 | 0.3619 | 0.2818 | 0.2944 | 0.3118 | 0.1971 | 0.2297 | 0.2635 | x | 0.3109 | 0.3098 | 0.475 | 0.405 | 0.3939 | 0.3379 | 0.4851 | 0.3984 | 0.434 | 0.3618 | 0.2824 | 0.2937 | 0.312 | 0.1985 | 0.2303 | 0.2637
y | 0.3307 | 0.326 | 0.3271 | 0.3264 | 0.3409 | 0.3252 | 0.4638 | 0.4117 | 0.4474 | 0.3836 | 0.402 | 0.3521 | 0.3609 | 0.254 | 0.2835 | 0.305 | y | 0.331 | 0.3266 | 0.3284 | 0.3289 | 0.3414 | 0.3253 | 0.4647 | 0.4148 | 0.4492 | 0.3841 | 0.4068 | 0.3552 | 0.3613 | 0.2536 | 0.2834 | 0.305 | y | 0.3307 | 0.3257 | 0.3274 | 0.3274 | 0.3413 | 0.3252 | 0.4638 | 0.4124 | 0.4469 | 0.3839 | 0.4041 | 0.3527 | 0.3612 | 0.2549 | 0.284 | 0.3052
Wcie | 87.13 | 22.46 | -113.5 | -57.71 | -52.66 | 47.43 | -310.2 | -191.3 | -266 | -55.92 | -83.06 | -11 | 11.98 | 233.54 | 172.16 | 129.78 | Wcie | 86.6 | 20.78 | -121.3 | -66.36 | -54.76 | 46.76 | -312.8 | -199.6 | -270.8 | -57.57 | -91.65 | -17.25 | 10.78 | 235.31 | 172.67 | 129.85 | Wcie | 87.03 | 23 | -114.3 | -60.42 | -54.41 | 47.25 | -310.9 | -193.9 | -265.4 | -57.33 | -87.42 | -12.46 | 10.58 | 232.13 | 171.24 | 129.19
Tcie | 2.53 | 0.27 | -147.6 | -84.03 | -64.8 | -25.35 | -67.23 | -22.35 | -32.23 | -8.4 | 75.17 | 32.08 | 21.75 | 53.63 | 44.3 | 28.26 | Tcie | 2.43 | -0.02 | -152.6 | -87.3 | -65.79 | -25.71 | -67.84 | -23.42 | -32.94 | -8.71 | 78.02 | 33.18 | 21.58 | 54.72 | 44.74 | 28.36 | Tcie | 2.43 | 0.2 | -147.3 | -84.36 | -65.43 | -25.46 | -67.82 | -23.14 | -32.85 | -8.81 | 75.79 | 32.24 | 21.23 | 54.35 | 44.65 | 28.33
Wgg | 96.06 | 41.56 | -273.5 | -140.2 | -152.4 | 38.25 | -754.3 | -437.1 | -623 | -207 | -185.7 | -26.45 | -40.33 | 516.23 | 361.83 | 240.73 | Wgg | 93.74 | 37.13 | -292 | -160.3 | -157.9 | 35.97 | -761.2 | -456.3 | -634.7 | -211.5 | -205.1 | -40.71 | -43.74 | 519.78 | 362.46 | 240.15 | Wgg
Tgg | 4.28 | 1.8 | -166.2 | -93.97 | -72.41 | -27.29 | -77.55 | -25.55 | -37.47 | -9.15 | 85.43 | 37.43 | 25.53 | 63.81 | 52.65 | 34.01 | Tgg | 3.95 | 1.26 | -172.2 | -98.02 | -73.83 | -27.92 | -78.75 | -27.2 | -38.74 | -9.84 | 88.3 | 38.41 | 25.08 | 65.02 | 53.05 | 33.97 | Tgg
L* | 93.5 | 39.43 | 35.48 | 30.88 | 59.44 | 80.07 | 77.23 | 41.63 | 59.73 | 87.65 | 40.99 | 35.14 | 82.3 | 38.89 | 56.74 | 73.15 | L* | 93.6 | 39.17 | 35.04 | 30.31 | 59.29 | 80.02 | 77.27 | 41.32 | 59.61 | 87.56 | 40.49 | 34.53 | 82.23 | 38.55 | 56.57 | 73.11 | L* | 93.54 | 39.19 | 35.03 | 30.56 | 59.09 | 79.94 | 77.15 | 41.45 | 59.55 | 87.45 | 40.71 | 34.86 | 82.1 | 38.8 | 56.67 | 73.09
a* | -1.38 | 0.21 | 33.89 | 18.75 | 21.89 | 12.79 | 13.18 | 1.39 | 2.42 | -1.05 | -23.58 | -9.28 | -13.17 | -14.71 | -15.75 | -11.66 | a* | -1.36 | 0.26 | 34.41 | 18.95 | 22.11 | 12.94 | 13.3 | 1.48 | 2.5 | -0.95 | -24.17 | -9.54 | -13.1 | -15.07 | -15.92 | -11.7 | a* | -1.33 | 0.25 | 33.5 | 18.59 | 21.95 | 12.84 | 13.36 | 1.53 | 2.63 | -0.88 | -23.63 | -9.29 | -12.91 | -15.01 | -15.92 | -11.7
b* | -0.66 | -1.27 | 15.43 | 6.75 | 13.09 | 1.92 | 85.12 | 24.17 | 48.59 | 26.31 | 9.64 | 1.86 | 9.17 | -24.75 | -21.21 | -14.83 | b* | -0.5 | -1.1 | 16.38 | 7.56 | 13.4 | 2.04 | 86.33 | 25.1 | 49.52 | 26.61 | 10.37 | 2.4 | 9.37 | -24.81 | -21.25 | -14.84 | b* | -0.62 | -1.34 | 15.35 | 6.97 | 13.26 | 1.91 | 85.37 | 24.43 | 48.37 | 26.48 | 10.02 | 1.97 | 9.35 | -24.52 | -21.05 | -14.73
Name: | 00 white | 01 gray | 02 bright red | 03 lidka red | 04 dark pink | 05 gosia pink | 06 yellow | 07 adam olive | 07 yellow olive | 09 yellow beige | 10 green | 11 forest green (dark) | 12 mint | 13 dark blue | 14 blue | 15 light blue | Name: | 00 white | 01 gray | 02 bright red | 03 lidka red | 04 dark pink | 05 gosia pink | 06 yellow | 07 adam olive | 08 yellow olive | 09 yellow beige | 10 green | 11 forest green (dark) | 12 mint | 13 dark blue | 14 blue | 15 light blue | Name: | 00 white | 01 gray | 02 bright red | 03 lidka red | 04 dark pink | 05 gosia pink | 06 yellow | 07 adam olive | 08 yellow olive | 09 yellow beige | 10 green | 11 forest green (dark) | 12 mint | 13 dark blue | 14 blue | 15 light blue
360 | 22.73 | 10.22 | 4.83 | 5.15 | 17.93 | 20.2 | 4.93 | 4.88 | 6.35 | 11.14 | 6.28 | 5.76 | 9.85 | 18.03 | 30.12 | 48.04 | 360 | [16 empty] | 360
370 | 35.08 | 10.72 | 5.15 | 5.55 | 19.05 | 26.31 | 5.32 | 5.23 | 6.84 | 13.21 | 6.82 | 6.41 | 11.96 | 20 | 32.8 | 52.26 | 370 | [16 empty] | 370
380 | 54.48 | 10.9 | 5.1 | 5.39 | 19.25 | 34.13 | 5.21 | 5.3 | 6.88 | 16.27 | 7.73 | 7.24 | 15.81 | 20.14 | 34.31 | 53.91 | 380 | [16 empty] | 380
390 | 69.45 | 11.27 | 5.09 | 5.38 | 19.32 | 42.62 | 5.15 | 5.31 | 6.76 | 21.23 | 8.38 | 7.76 | 22.86 | 19.54 | 35.32 | 55.33 | 390 | [16 empty] | 390
400 | 77.96 | 11.53 | 5.11 | 5.33 | 19.62 | 50.91 | 5.09 | 5.26 | 6.73 | 26.85 | 8.53 | 7.92 | 31.81 | 18.67 | 35.69 | 56.03 | 400 | 78.01 | 11.24 | 4.65 | 4.84 | 19.44 | 50.98 | 4.82 | 4.87 | 6.39 | 26.82 | 8 | 7.38 | 31.81 | 18.46 | 35.67 | 56.24 | 400 | 77.86 | 11.47 | 4.86 | 5 | 19.37 | 50.62 | 4.98 | 5.07 | 6.65 | 26.78 | 8.23 | 7.78 | 31.55 | 18.61 | 35.67 | 56.14
410 | 81.6 | 11.61 | 4.99 | 5.23 | 19.83 | 55.38 | 5.01 | 5.2 | 6.72 | 31.03 | 8.46 | 7.91 | 38.47 | 18.01 | 35.81 | 56.51 | 410 | 81.69 | 11.41 | 4.65 | 4.77 | 19.53 | 55.19 | 4.78 | 4.87 | 6.35 | 30.67 | 7.93 | 7.44 | 38.1 | 17.73 | 35.63 | 56.62 | 410 | 81.66 | 11.56 | 4.85 | 5.02 | 19.39 | 54.99 | 4.93 | 5.11 | 6.62 | 30.69 | 8.2 | 7.78 | 37.98 | 17.93 | 35.7 | 56.5
420 | 83.12 | 11.61 | 4.97 | 5.2 | 19.82 | 56.5 | 4.94 | 5.17 | 6.64 | 33.73 | 8.32 | 7.87 | 42.3 | 18.46 | 36.39 | 57.33 | 420 | 83.11 | 11.35 | 4.63 | 4.76 | 19.52 | 56.23 | 4.7 | 4.86 | 6.27 | 33.33 | 7.81 | 7.37 | 41.87 | 18.16 | 36.2 | 57.24 | 420 | 83.23 | 11.49 | 4.81 | 4.99 | 19.4 | 56.26 | 4.84 | 5.07 | 6.56 | 33.41 | 8.04 | 7.67 | 41.8 | 18.3 | 36.23 | 57.17
430 | 84.16 | 11.55 | 4.97 | 5.19 | 19.86 | 56.73 | 4.98 | 5.18 | 6.6 | 36.48 | 8.35 | 7.93 | 45.37 | 19.38 | 37.5 | 58.29 | 430 | 84.12 | 11.31 | 4.6 | 4.72 | 19.58 | 56.41 | 4.65 | 4.84 | 6.25 | 36.08 | 7.85 | 7.43 | 44.98 | 19.11 | 37.29 | 58.21 | 430 | 84.18 | 11.44 | 4.79 | 4.93 | 19.43 | 56.42 | 4.81 | 5.05 | 6.53 | 36.09 | 8.06 | 7.7 | 44.84 | 19.22 | 37.27 | 58.06
440 | 84.77 | 11.43 | 4.92 | 5.08 | 19.89 | 56.39 | 4.93 | 5.16 | 6.51 | 39.52 | 8.4 | 7.94 | 48.2 | 20.05 | 38.38 | 58.97 | 440 | 84.66 | 11.22 | 4.58 | 4.71 | 19.6 | 56.17 | 4.66 | 4.84 | 6.24 | 39.1 | 7.92 | 7.48 | 47.87 | 19.77 | 38.11 | 58.85 | 440 | 84.81 | 11.33 | 4.79 | 4.93 | 19.48 | 56.16 | 4.82 | 5.05 | 6.53 | 39.11 | 8.16 | 7.78 | 47.72 | 19.92 | 38.16 | 58.69
450 | 85.24 | 11.36 | 4.92 | 5.05 | 19.94 | 55.81 | 4.9 | 5.16 | 6.51 | 42.87 | 8.45 | 7.95 | 50.95 | 21 | 39.39 | 59.76 | 450 | 85.24 | 11.16 | 4.57 | 4.73 | 19.67 | 55.58 | 4.66 | 4.88 | 6.25 | 42.53 | 7.99 | 7.54 | 50.66 | 20.73 | 39.18 | 59.72 | 450 | 85.3 | 11.25 | 4.78 | 4.93 | 19.54 | 55.57 | 4.8 | 5.04 | 6.49 | 42.46 | 8.2 | 7.79 | 50.5 | 20.81 | 39.16 | 59.56
460 | 85.69 | 11.3 | 4.9 | 5.05 | 20.02 | 54.77 | 4.89 | 5.17 | 6.5 | 46.66 | 8.47 | 7.97 | 53.82 | 22.47 | 40.7 | 60.88 | 460 | 85.79 | 11.07 | 4.57 | 4.71 | 19.73 | 54.63 | 4.65 | 4.88 | 6.23 | 46.24 | 8.02 | 7.57 | 53.55 | 22.13 | 40.5 | 60.83 | 460 | 85.77 | 11.16 | 4.77 | 4.91 | 19.64 | 54.61 | 4.81 | 5.04 | 6.48 | 46.15 | 8.22 | 7.83 | 53.33 | 22.21 | 40.43 | 60.64
470 | 85.92 | 11.21 | 4.9 | 5.04 | 20.1 | 53.52 | 4.9 | 5.18 | 6.48 | 50.44 | 8.5 | 7.99 | 56.59 | 23.64 | 41.6 | 61.68 | 470 | 85.98 | 11 | 4.57 | 4.7 | 19.83 | 53.33 | 4.67 | 4.88 | 6.21 | 50.08 | 8.07 | 7.58 | 56.34 | 23.38 | 41.43 | 61.7 | 470 | 85.93 | 11.07 | 4.77 | 4.89 | 19.74 | 53.3 | 4.85 | 5.06 | 6.47 | 49.95 | 8.28 | 7.84 | 56.03 | 23.39 | 41.36 | 61.48
480 | 86.08 | 11.15 | 4.93 | 5.04 | 20.24 | 51.93 | 4.97 | 5.24 | 6.53 | 54.56 | 8.63 | 8.06 | 59.52 | 23.91 | 41.8 | 61.95 | 480 | 86.01 | 10.94 | 4.6 | 4.72 | 19.94 | 51.79 | 4.77 | 4.93 | 6.27 | 54.04 | 8.21 | 7.62 | 59.13 | 23.65 | 41.57 | 61.87 | 480 | 86.04 | 11.01 | 4.78 | 4.9 | 19.85 | 51.83 | 4.95 | 5.1 | 6.53 | 53.95 | 8.42 | 7.88 | 58.85 | 23.7 | 41.56 | 61.74
490 | 85.82 | 11.08 | 4.93 | 5.01 | 20.32 | 50.39 | 5.22 | 5.36 | 6.79 | 57.94 | 8.99 | 8.19 | 61.8 | 23.93 | 41.63 | 62.01 | 490 | 85.98 | 10.89 | 4.62 | 4.72 | 20.08 | 50.2 | 5.08 | 5.08 | 6.57 | 57.76 | 8.64 | 7.81 | 61.67 | 23.68 | 41.45 | 61.88 | 490 | 85.91 | 10.94 | 4.79 | 4.89 | 19.95 | 50.26 | 5.28 | 5.26 | 6.84 | 57.55 | 8.86 | 8.06 | 61.33 | 23.7 | 41.43 | 61.75
500 | 85.55 | 11.03 | 4.96 | 5 | 20.44 | 48.96 | 6.7 | 6.25 | 8.55 | 61.42 | 10.98 | 9.01 | 64.06 | 24.27 | 41.71 | 62.14 | 500 | 85.73 | 10.84 | 4.67 | 4.74 | 20.21 | 48.77 | 6.61 | 5.96 | 8.33 | 61.18 | 10.6 | 8.64 | 63.85 | 24.16 | 41.61 | 62.11 | 500 | 85.63 | 10.89 | 4.83 | 4.89 | 20.1 | 48.81 | 7.06 | 6.21 | 8.78 | 60.93 | 10.9 | 8.86 | 63.41 | 23.93 | 41.46 | 61.81
510 | 85.28 | 10.97 | 5.01 | 5 | 20.63 | 48.05 | 16.18 | 9.98 | 17.98 | 64.63 | 17.45 | 10.89 | 65.8 | 21.67 | 39.81 | 60.55 | 510 | 85.39 | 10.8 | 4.74 | 4.74 | 20.41 | 47.92 | 16.17 | 9.72 | 17.73 | 64.28 | 17 | 10.52 | 65.53 | 21.49 | 39.6 | 60.49 | 510 | 85.26 | 10.84 | 4.88 | 4.89 | 20.3 | 47.92 | 16.24 | 9.64 | 17.41 | 64.02 | 16.79 | 10.56 | 65.1 | 21.52 | 39.6 | 60.31
520 | 84.88 | 10.95 | 5.09 | 5.01 | 21.08 | 48.1 | 34.48 | 12.51 | 27.7 | 67.35 | 21.35 | 11.27 | 66.71 | 16.46 | 35.18 | 56.72 | 520 | 84.95 | 10.76 | 4.82 | 4.75 | 20.83 | 47.85 | 34.3 | 12.29 | 27.53 | 67.05 | 21.12 | 10.97 | 66.44 | 16.31 | 35.06 | 56.68 | 520 | 84.88 | 10.79 | 4.94 | 4.88 | 20.7 | 47.84 | 33.89 | 12.27 | 27.08 | 66.8 | 20.82 | 11.05 | 66.07 | 16.53 | 35.15 | 56.58
530 | 84.53 | 10.9 | 5.19 | 5.03 | 21.69 | 49.01 | 49.33 | 12.99 | 30.67 | 69.79 | 20.03 | 10.72 | 66.88 | 11.64 | 29.63 | 51.72 | 530 | 84.78 | 10.75 | 4.95 | 4.8 | 21.5 | 48.84 | 49.22 | 12.8 | 30.48 | 69.53 | 19.77 | 10.45 | 66.7 | 11.51 | 29.54 | 51.71 | 530 | 84.66 | 10.76 | 5.06 | 4.92 | 21.34 | 48.8 | 48.84 | 12.87 | 30.36 | 69.34 | 19.81 | 10.58 | 66.37 | 11.78 | 29.68 | 51.69
540 | 84.57 | 10.87 | 5.31 | 5.05 | 22.28 | 50.89 | 58.02 | 13.19 | 31.52 | 71.8 | 16.71 | 9.97 | 66.31 | 8.42 | 24.64 | 46.65 | 540 | 84.83 | 10.73 | 5.08 | 4.85 | 22.12 | 50.8 | 58.16 | 13 | 31.41 | 71.68 | 16.41 | 9.69 | 66.17 | 8.22 | 24.47 | 46.63 | 540 | 84.7 | 10.75 | 5.19 | 4.95 | 21.94 | 50.69 | 57.79 | 13.12 | 31.34 | 71.44 | 16.64 | 9.86 | 65.89 | 8.51 | 24.69 | 46.68
550 | 84.81 | 10.88 | 5.43 | 5.08 | 22.76 | 53.55 | 63.35 | 13.4 | 32.13 | 73.69 | 13.46 | 9.18 | 65.14 | 6.5 | 20.58 | 42.1 | 550 | 85.08 | 10.73 | 5.18 | 4.87 | 22.6 | 53.45 | 63.54 | 13.22 | 32.04 | 73.56 | 13.12 | 8.88 | 65.01 | 6.26 | 20.37 | 42.02 | 550 | 84.87 | 10.74 | 5.29 | 4.98 | 22.42 | 53.28 | 63.17 | 13.32 | 31.95 | 73.3 | 13.39 | 9.06 | 64.73 | 6.53 | 20.62 | 42.15
560 | 84.65 | 10.89 | 5.49 | 5.09 | 23.27 | 56.62 | 66.91 | 13.6 | 32.65 | 75.37 | 10.53 | 8.34 | 63.39 | 5.66 | 18.21 | 39.29 | 560 | 84.96 | 10.73 | 5.27 | 4.89 | 23.09 | 56.42 | 66.95 | 13.41 | 32.52 | 75.22 | 10.28 | 8.06 | 63.34 | 5.48 | 18.05 | 39.26 | 560 | 84.75 | 10.73 | 5.38 | 5 | 22.96 | 56.24 | 66.65 | 13.5 | 32.47 | 74.95 | 10.55 | 8.24 | 63.09 | 5.66 | 18.21 | 39.29
570 | 83.94 | 10.87 | 5.6 | 5.13 | 24.6 | 59.58 | 69.18 | 13.78 | 33.1 | 76.82 | 8.61 | 7.61 | 61.46 | 5.34 | 16.97 | 37.66 | 570 | 84.23 | 10.71 | 5.4 | 4.94 | 24.43 | 59.38 | 69.26 | 13.59 | 32.95 | 76.63 | 8.34 | 7.34 | 61.32 | 5.14 | 16.78 | 37.56 | 570 | 84.09 | 10.7 | 5.51 | 5.04 | 24.29 | 59.17 | 68.96 | 13.66 | 32.87 | 76.36 | 8.57 | 7.52 | 61.12 | 5.29 | 16.88 | 37.55
580 | 82.95 | 10.82 | 5.99 | 5.3 | 28.07 | 62.24 | 70.7 | 13.93 | 33.42 | 77.99 | 7.75 | 7.15 | 59.24 | 5.04 | 15.62 | 35.69 | 580 | 83.24 | 10.7 | 5.83 | 5.14 | 27.94 | 62.19 | 70.88 | 13.77 | 33.33 | 77.84 | 7.47 | 6.88 | 59.15 | 4.84 | 15.43 | 35.62 | 580 | 83.14 | 10.68 | 5.92 | 5.24 | 27.65 | 61.94 | 70.6 | 13.85 | 33.24 | 77.56 | 7.64 | 7.04 | 58.96 | 4.98 | 15.57 | 35.65
590 | 82.26 | 10.79 | 7.34 | 6.07 | 33.29 | 64.68 | 71.86 | 14.11 | 33.77 | 78.93 | 7.3 | 6.86 | 57 | 4.84 | 14.52 | 34.09 | 590 | 82.58 | 10.68 | 7.22 | 5.87 | 33.2 | 64.74 | 72.12 | 13.97 | 33.7 | 78.85 | 7.01 | 6.59 | 56.94 | 4.64 | 14.31 | 33.99 | 590 | 82.44 | 10.67 | 7.19 | 5.99 | 32.81 | 64.48 | 71.83 | 14.04 | 33.59 | 78.57 | 7.17 | 6.74 | 56.78 | 4.77 | 14.46 | 34.02
600 | 82.14 | 10.79 | 10.72 | 8.15 | 38.88 | 66.97 | 72.71 | 14.32 | 34.08 | 79.51 | 6.89 | 6.66 | 54.74 | 4.83 | 14.12 | 33.52 | 600 | 82.47 | 10.68 | 10.72 | 7.9 | 38.87 | 67.02 | 73.01 | 14.18 | 34.06 | 79.5 | 6.59 | 6.39 | 54.72 | 4.63 | 13.91 | 33.4 | 600 | 82.31 | 10.67 | 10.39 | 8 | 38.41 | 66.76 | 72.75 | 14.25 | 33.94 | 79.3 | 6.76 | 6.54 | 54.62 | 4.76 | 14.04 | 33.44
610 | 82.17 | 10.79 | 16.84 | 11.68 | 44.67 | 68.94 | 73.41 | 14.54 | 34.39 | 79.81 | 6.59 | 6.53 | 52.47 | 4.97 | 14.32 | 33.81 | 610 | 82.49 | 10.67 | 16.98 | 11.35 | 44.64 | 69.04 | 73.68 | 14.4 | 34.37 | 79.83 | 6.3 | 6.26 | 52.51 | 4.77 | 14.13 | 33.77 | 610 | 82.34 | 10.65 | 16.18 | 11.32 | 44.1 | 68.75 | 73.4 | 14.45 | 34.25 | 79.64 | 6.46 | 6.4 | 52.47 | 4.88 | 14.19 | 33.71
620 | 82.13 | 10.8 | 25.69 | 15.2 | 50.26 | 70.7 | 73.92 | 14.81 | 34.76 | 79.85 | 6.56 | 6.53 | 50.41 | 5.06 | 14.48 | 34.01 | 620 | 82.44 | 10.68 | 25.4 | 14.85 | 50.22 | 70.79 | 74.2 | 14.64 | 34.69 | 79.91 | 6.25 | 6.25 | 50.46 | 4.87 | 14.31 | 33.98 | 620 | 82.27 | 10.64 | 24.49 | 14.77 | 49.63 | 70.49 | 73.95 | 14.69 | 34.54 | 79.69 | 6.36 | 6.38 | 50.42 | 4.96 | 14.33 | 33.91
630 | 81.99 | 10.81 | 35.64 | 17.36 | 55.15 | 72.12 | 74.22 | 15.04 | 35.03 | 79.67 | 6.58 | 6.54 | 48.45 | 5.05 | 14.26 | 33.62 | 630 | 82.3 | 10.7 | 35.11 | 17.07 | 55.19 | 72.27 | 74.56 | 14.9 | 34.99 | 79.78 | 6.3 | 6.3 | 48.51 | 4.87 | 14.07 | 33.57 | 630 | 82.16 | 10.66 | 34.4 | 17.05 | 54.67 | 72.01 | 74.33 | 14.93 | 34.85 | 79.6 | 6.4 | 6.41 | 48.49 | 4.97 | 14.17 | 33.57
640 | 81.9 | 10.83 | 45.79 | 18.6 | 59.17 | 73.43 | 74.51 | 15.29 | 35.34 | 79.55 | 6.56 | 6.55 | 46.67 | 5.14 | 14.2 | 33.57 | 640 | 82.16 | 10.72 | 45.14 | 18.38 | 59.26 | 73.57 | 74.82 | 15.15 | 35.31 | 79.62 | 6.29 | 6.3 | 46.71 | 4.94 | 14 | 33.47 | 640 | 82.05 | 10.71 | 44.67 | 18.42 | 58.8 | 73.33 | 74.57 | 15.2 | 35.17 | 79.45 | 6.4 | 6.42 | 46.71 | 5.05 | 14.12 | 33.49
650 | 82.08 | 10.87 | 54.48 | 20.37 | 62.23 | 74.57 | 74.61 | 15.52 | 35.59 | 79.29 | 6.48 | 6.55 | 44.99 | 5.53 | 14.93 | 34.88 | 650 | 82.39 | 10.76 | 53.93 | 20.1 | 62.43 | 74.76 | 75.03 | 15.4 | 35.6 | 79.46 | 6.21 | 6.32 | 45.08 | 5.34 | 14.7 | 34.72 | 650 | 82.26 | 10.75 | 53.7 | 20.14 | 62 | 74.53 | 74.79 | 15.47 | 35.47 | 79.31 | 6.34 | 6.43 | 45.11 | 5.44 | 14.8 | 34.74
660 | 82.76 | 10.88 | 60.76 | 24.05 | 64.56 | 75.58 | 74.66 | 15.72 | 35.85 | 79.12 | 6.45 | 6.65 | 43.48 | 6.72 | 17.24 | 38.69 | 660 | 83.06 | 10.8 | 60.48 | 23.72 | 64.79 | 75.82 | 75.07 | 15.63 | 35.86 | 79.29 | 6.19 | 6.43 | 43.56 | 6.53 | 16.95 | 38.46 | 660 | 82.93 | 10.77 | 60.29 | 23.63 | 64.39 | 75.58 | 74.88 | 15.68 | 35.75 | 79.16 | 6.31 | 6.53 | 43.61 | 6.59 | 16.96 | 38.28
670 | 83.89 | 10.87 | 64.85 | 30.48 | 66.29 | 76.52 | 74.68 | 15.89 | 36.06 | 78.95 | 6.65 | 6.97 | 42.12 | 9.74 | 21.9 | 45.36 | 670 | 84.17 | 10.81 | 64.81 | 30.09 | 66.53 | 76.72 | 74.99 | 15.8 | 36.05 | 79.13 | 6.4 | 6.75 | 42.14 | 9.4 | 21.49 | 45 | 670 | 83.96 | 10.76 | 64.55 | 29.75 | 66.09 | 76.48 | 74.78 | 15.85 | 35.89 | 78.93 | 6.49 | 6.82 | 42.18 | 9.38 | 21.38 | 44.61
680 | 85.53 | 10.87 | 67.61 | 39.75 | 67.73 | 77.56 | 74.62 | 16.05 | 36.22 | 78.82 | 7.53 | 7.84 | 40.8 | 15.91 | 29.69 | 54.36 | 680 | 85.79 | 10.8 | 67.55 | 39.19 | 67.82 | 77.67 | 74.87 | 15.94 | 36.19 | 78.88 | 7.28 | 7.58 | 40.84 | 15.39 | 29.06 | 53.7 | 680 | 85.55 | 10.76 | 67.3 | 38.47 | 67.46 | 77.44 | 74.73 | 16.03 | 36.06 | 78.75 | 7.33 | 7.62 | 40.93 | 15.09 | 28.65 | 53.07
690 | 86.98 | 10.75 | 68.64 | 48.13 | 68.32 | 78.05 | 74.14 | 16 | 35.97 | 78.21 | 9.97 | 9.39 | 39.43 | 24.32 | 38.51 | 61.72 | 690 | 87.57 | 10.8 | 69.25 | 48.71 | 68.95 | 78.61 | 74.82 | 16.1 | 36.36 | 78.64 | 9.8 | 9.28 | 39.79 | 24.45 | 38.79 | 62.18 | 690 | 87.21 | 10.74 | 68.96 | 47.95 | 68.52 | 78.34 | 74.55 | 16.2 | 36.2 | 78.41 | 9.82 | 9.27 | 39.85 | 23.99 | 38.28 | 61.51
700 | 88.25 | 10.74 | 69.5 | 55.45 | 69.18 | 78.69 | 73.87 | 16.13 | 36.08 | 77.81 | 14.54 | 11.42 | 38.55 | 33.75 | 47.75 | 67.75 | 700 | 89.08 | 10.77 | 70.23 | 57.12 | 69.92 | 79.4 | 74.55 | 16.23 | 36.43 | 78.17 | 15.67 | 11.69 | 38.79 | 35.62 | 49.55 | 68.86 | 700 | 88.75 | 10.71 | 69.96 | 56.33 | 69.48 | 79.13 | 74.39 | 16.35 | 36.33 | 77.99 | 15.35 | 11.62 | 38.85 | 34.85 | 48.74 | 68.25

## GROUP 3 SAMPLES
DATA FOR DATACOLOR SPECTRAFLASH 450 | [24 empty] | DATA FOR DATACOLOR SPECTRAFLASH 600 | [24 empty] | DATA FOR DATACOLOR SPECTRAFLASH 300
Inst, no: | 435 | 435 | 435 | 435 | 435 | 435 | 435 | 435 | 435 | 435 | 435 | 435 | 435 | 435 | 435 | 435 | 435 | 435 | 435 | 435 | 435 | 435 | 435 | 435 | Inst, no: | 596 | 596 | 596 | 596 | 596 | 596 | 596 | 596 | 596 | 596 | 596 | 596 | 596 | 596 | 596 | 596 | 596 | 596 | 596 | 596 | 596 | 596 | 596 | 596 | Inst, no: | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118
Conditions: | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | Conditions: | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | LAV I Calib | Conditions: | %R LAV SCI Calib | %R LAV SCI Calib | %R LAV SCI Calib | LAV I Calib | %R LAV SCI Calib | %R LAV SCI Calib | %R LAV SCI Calib | %R LAV SCI Calib | %R LAV SCI Calib | %R LAV SCI Calib | LAV I Calib | %R LAV SCI Calib | %R LAV SCI Calib | %R LAV SCI Calib | %R LAV SCI Calib | %R LAV SCI Calib | %R LAV SCI Calib | %R LAV SCI Calib | LAV I Calib | %R LAV SCI Calib | %R LAV SCI Calib | %R LAV SCI Calib | %R LAV SCI Calib | %R LAV SCI Calib
Name: | cc1 | cc2 | cc3 | cc4 | cc5 | cc6 | cc7 | cc8 | cc9 | cc10 | cc11 | cc12 | cc13 | cc14 | cc15 | cc16 | cc17 | cc18 | cc19 | cc20 | cc21 | cc22 | cc23 | cc24 | Name: | CC1 | CC2 | CC3 | CC4 | CC5 | CC6 | CC7 | CC8 | CC9 | CC10 | CC11 | CC12 | CC13 | CC14 | CC15 | CC16 | CC17 | CC18 | CC19 | CC20 | CC21 | CC22 | CC23 | CC24 | Name: | CC1 | CC2 | CC3 | CC4 | CC5 | CC6 | CC7 | CC8 | CC9 | CC10 | CC11 | CC12 | CC13 | CC14 | CC15 | CC16 | CC17 | CC18 | CC19 | CC20 | CC21 | CC22 | CC23 | CC24
X | 11.11 | 35.89 | 17.43 | 10.7 | 24.49 | 31.05 | 36.07 | 13.68 | 26.55 | 8.76 | 33.46 | 43.04 | 8.15 | 15 | 17.9 | 53.71 | 28.82 | 14.69 | 83.27 | 53.78 | 33.71 | 17.81 | 8.82 | 3.65 | X | 11.24 | 3.68 | 19.14 | 10.73 | 24.7 | 31.25 | 36.37 | 13.76 | 26.75 | 8.81 | 33.79 | 43.45 | 8.23 | 15.09 | 18.02 | 54.23 | 28.99 | 14.84 | 83.89 | 54.28 | 33.89 | 17.93 | 8.75 | 3.65 | X | 10.62 | 34.73 | 16.93 | 10.39 | 23.86 | 30.34 | 35.26 | 13.3 | 25.89 | 8.44 | 32.82 | 41.88 | 7.87 | 14.69 | 17.36 | 52.62 | 27.95 | 14.22 | 80.91 | 52.61 | 33.13 | 17.36 | 8.54 | 3.55
Y | 10.05 | 33.78 | 19.25 | 12.9 | 24.28 | 42.18 | 28.15 | 13.17 | 18.54 | 7.24 | 41.69 | 38.87 | 7.56 | 22.68 | 11.3 | 54.92 | 20.14 | 20.83 | 88.46 | 56.96 | 35.69 | 18.83 | 9.34 | 3.85 | Y | 10.22 | 3.88 | 21.01 | 12.95 | 24.5 | 42.45 | 28.38 | 13.27 | 18.68 | 7.27 | 42.05 | 39.2 | 7.66 | 22.81 | 11.38 | 55.38 | 20.26 | 21.03 | 89.12 | 57.46 | 35.86 | 18.94 | 9.25 | 3.85 | Y | 9.6 | 32.66 | 18.7 | 12.51 | 23.69 | 41.17 | 27.46 | 12.9 | 18.07 | 6.97 | 40.69 | 37.74 | 7.38 | 22.09 | 10.97 | 53.62 | 19.56 | 20.19 | 85.9 | 55.69 | 35.05 | 18.33 | 9.03 | 3.75
Z | 6.84 | 24.16 | 33.32 | 6.87 | 43.01 | 42.93 | 6.31 | 39.08 | 13.66 | 14.87 | 10.57 | 7.41 | 27.81 | 9.51 | 5.34 | 8.33 | 31.06 | 38 | 89.65 | 60.66 | 38.58 | 20.2 | 10.45 | 4.28 | Z | 7.15 | 4.34 | 33.75 | 6.89 | 43.48 | 43.1 | 6.45 | 39.37 | 13.79 | 15.06 | 10.65 | 7.58 | 28.14 | 9.51 | 5.43 | 8.45 | 31.38 | 38.24 | 90.06 | 61.25 | 38.79 | 20.34 | 10.36 | 4.31 | Z | 6.5 | 23.21 | 32.25 | 6.6 | 41.9 | 41.59 | 6.12 | 37.94 | 13.27 | 14.42 | 10.24 | 7.19 | 26.88 | 9.23 | 5.17 | 8.11 | 30.21 | 36.45 | 86.2 | 59.11 | 37.78 | 19.62 | 10.08 | 4.17
x | 0.3967 | 0.3824 | 0.249 | 0.3511 | 0.2668 | 0.2673 | 0.5113 | 0.2075 | 0.4519 | 0.2838 | 0.3904 | 0.4819 | 0.1873 | 0.3179 | 0.5181 | 0.4593 | 0.3601 | 0.1998 | 0.3186 | 0.3138 | 0.3122 | 0.3134 | 0.3084 | 0.3095 | x | 0.3928 | 0.3091 | 0.259 | 0.351 | 0.2665 | 0.2676 | 0.5108 | 0.2072 | 0.4517 | 0.2828 | 0.3907 | 0.4816 | 0.187 | 0.3184 | 0.5173 | 0.4593 | 0.3595 | 0.2002 | 0.3189 | 0.3138 | 0.3122 | 0.3134 | 0.3085 | 0.3092 | x | 0.3976 | 0.3833 | 0.2494 | 0.3522 | 0.2667 | 0.2682 | 0.5122 | 0.2073 | 0.4523 | 0.283 | 0.3919 | 0.4824 | 0.1869 | 0.3192 | 0.5183 | 0.4602 | 0.3596 | 0.2007 | 0.3198 | 0.3143 | 0.3127 | 0.3139 | 0.3089 | 0.3096
y | 0.359 | 0.36 | 0.275 | 0.4234 | 0.2645 | 0.3631 | 0.3992 | 0.1998 | 0.3156 | 0.2345 | 0.4863 | 0.4352 | 0.1738 | 0.4805 | 0.3273 | 0.4695 | 0.2517 | 0.2834 | 0.3384 | 0.3323 | 0.3305 | 0.3313 | 0.3264 | 0.3267 | y | 0.3572 | 0.3261 | 0.2843 | 0.4235 | 0.2643 | 0.3634 | 0.3986 | 0.1998 | 0.3154 | 0.2335 | 0.4861 | 0.4345 | 0.1739 | 0.4811 | 0.3268 | 0.4691 | 0.2512 | 0.2838 | 0.3388 | 0.3322 | 0.3304 | 0.3311 | 0.3262 | 0.3261 | y | 0.3592 | 0.3605 | 0.2755 | 0.4241 | 0.2648 | 0.364 | 0.3989 | 0.2011 | 0.3158 | 0.2336 | 0.4858 | 0.4348 | 0.1751 | 0.4801 | 0.3275 | 0.4689 | 0.2517 | 0.2849 | 0.3395 | 0.3327 | 0.3308 | 0.3315 | 0.3267 | 0.3265
Wcie | -103.9 | -70.46 | 166.27 | -174 | 174.89 | 24.8 | -245.7 | 321.23 | -65.69 | 195.28 | -283.6 | -272.6 | 376.12 | -234.7 | -145.8 | -296.9 | 117.88 | 193.05 | 72.01 | 54.75 | 37.81 | 18.73 | 21.54 | 14.55 | Wcie | -97.47 | 16.08 | 144.28 | -174.1 | 175.67 | 24.36 | -244 | 321.54 | -65.05 | 197.75 | -283.1 | -270.8 | 376.16 | -236 | -144.3 | -295.7 | 119.29 | 192.16 | 71.87 | 55.53 | 38.16 | 19.18 | 21.61 | 15.92 | Wcie | -105.3 | -73.1 | 164.56 | -176.4 | 173.84 | 21.49 | -246.6 | 318.88 | -66.9 | 197.26 | -285 | -273.5 | 373.96 | -235.7 | -146.6 | -297.8 | 117.79 | 189.05 | 66.65 | 52.53 | 36.32 | 17.49 | 20.32 | 14.69
Tcie | -56.37 | -42.89 | 21.91 | 26.52 | -0.88 | 62.73 | -133.4 | 10.39 | -134.2 | -35.71 | 32.08 | -83.59 | 11.68 | 93.56 | -186.3 | -40.85 | -93.23 | 71.65 | 0.56 | 0.91 | 1.2 | 0.59 | 1.82 | 1.08 | Tcie | -54.1 | 1.09 | 18.94 | 26.64 | -0.73 | 62.67 | -133.3 | 10.7 | -134.2 | -35.43 | 31.67 | -83.71 | 12.04 | 93.45 | -185.8 | -41.21 | -92.98 | 71.57 | 0.48 | 0.79 | 1.04 | 0.44 | 1.65 | 1.01 | Tcie | -57.09 | -43.38 | 21.9 | 25.95 | -0.6 | 62.48 | -134.4 | 11.44 | -134.5 | -35.63 | 30.33 | -84.26 | 12.94 | 92.1 | -186.3 | -42.07 | -92.79 | 71.9 | 0.14 | 0.68 | 0.92 | 0.26 | 1.58 | 0.86
Wgg | -244 | -197.2 | 352.93 | -394.1 | 363.5 | 15.32 | -589.1 | 703.09 | -177.8 | 425.8 | -672 | -656.7 | 831.93 | -533.7 | -352.4 | -725.6 | 232.08 | 415.8 | 56.99 | 57.13 | 45.5 | 23.64 | 41.6 | 32.71 | Wgg | -230.5 | 35.21 | 300.88 | -395.3 | 364.2 | 13.13 | -586.7 | 703.04 | -177.5 | 430.53 | -672.5 | -654.1 | 831.31 | -537.7 | -350.2 | -724.6 | 234.28 | 412.8 | 55 | 57.39 | 45.2 | 23.63 | 41 | 34.87 | Wgg
Tgg | -63.18 | -47.89 | 27.37 | 29.74 | 1.68 | 72.04 | -151.5 | 15.74 | -150.8 | -37.3 | 34.85 | -95.58 | 17.71 | 104.79 | -210.2 | -47.68 | -102.9 | 83.72 | 1.9 | 2.4 | 2.78 | 2.06 | 3.55 | 2.71 | Tgg | -60.86 | 2.52 | 23.69 | 29.53 | 1.75 | 71.76 | -151.8 | 16.11 | -151.1 | -37.05 | 33.94 | -96.16 | 18.17 | 104.28 | -210 | -48.56 | -102.8 | 83.55 | 1.57 | 2.06 | 2.38 | 1.68 | 3.16 | 2.44 | Tgg
L* | 37.94 | 64.79 | 50.98 | 42.62 | 56.37 | 71 | 60.03 | 43.03 | 50.15 | 32.35 | 70.66 | 68.65 | 33.05 | 54.74 | 40.09 | 78.99 | 52 | 52.77 | 95.35 | 80.16 | 66.28 | 50.49 | 36.62 | 23.16 | L* | 38.23 | 23.28 | 52.96 | 42.68 | 56.58 | 71.18 | 60.23 | 43.17 | 50.31 | 32.42 | 70.9 | 68.89 | 33.26 | 54.87 | 40.22 | 79.26 | 52.13 | 52.99 | 95.63 | 80.44 | 66.41 | 50.62 | 36.47 | 23.19 | L* | 37.11 | 63.88 | 50.34 | 42.02 | 55.78 | 70.3 | 59.4 | 42.61 | 49.58 | 31.73 | 69.95 | 67.83 | 32.65 | 54.13 | 39.53 | 78.24 | 51.33 | 52.05 | 94.27 | 79.44 | 65.79 | 49.9 | 36.05 | 22.82
a* | 12.17 | 13.45 | -4.38 | -11.04 | 6.5 | -30.33 | 34.58 | 7.85 | 42 | 17.67 | -20.17 | 19.4 | 9.24 | -34.48 | 45.06 | 4.28 | 43.09 | -27.86 | -1.15 | -0.57 | -0.47 | -0.22 | -0.24 | -0.04 | a* | 11.86 | -0.02 | -3.9 | -11.08 | 6.47 | -30.36 | 34.72 | 7.72 | 42.14 | 17.73 | -20.08 | 19.56 | 9.09 | -34.5 | 45.14 | 4.44 | 43.19 | -27.9 | -1.14 | -0.5 | -0.39 | -0.15 | -0.18 | 0 | a* | 12.13 | 13.42 | -4.37 | -10.8 | 6.28 | -29.97 | 34.58 | 7.14 | 41.68 | 17.54 | -19.41 | 19.44 | 8.43 | -33.73 | 44.58 | 4.69 | 42.53 | -27.68 | -1.03 | -0.49 | -0.37 | -0.12 | -0.19 | 0.01
b* | 13.11 | 17.61 | -19.95 | 21.05 | -22.7 | 2.62 | 53.3 | -41.06 | 13.43 | -20.14 | 57.04 | 63.92 | -42.94 | 32.78 | 23.14 | 78.48 | -15.06 | -22.94 | 3.62 | 0.42 | -0.35 | 0.01 | -1.28 | -0.83 | b* | 12.41 | -0.95 | -17.12 | 21.08 | -22.85 | 2.74 | 53.07 | -41.17 | 13.38 | -20.45 | 57.22 | 63.7 | -43.08 | 33.04 | 22.96 | 78.51 | -15.3 | -22.86 | 3.81 | 0.37 | -0.38 | -0.03 | -1.29 | -0.94 | b* | 13.03 | 17.68 | -19.59 | 21.08 | -22.43 | 2.96 | 53.01 | -40.37 | 13.44 | -20.15 | 56.81 | 63.3 | -42.2 | 32.61 | 22.99 | 77.91 | -14.99 | -22.22 | 4.2 | 0.6 | -0.21 | 0.11 | -1.17 | -0.85
Name: | cc1 | cc2 | cc3 | cc4 | cc5 | cc6 | cc7 | cc8 | cc9 | cc10 | cc11 | cc12 | cc13 | cc14 | cc15 | cc16 | cc17 | cc18 | cc19 | cc20 | cc21 | cc22 | cc23 | cc24 | Name: | CC1 | CC2 | CC3 | CC4 | CC5 | CC6 | CC7 | CC8 | CC9 | CC10 | CC11 | CC12 | CC13 | CC14 | CC15 | CC16 | CC17 | CC18 | CC19 | CC20 | CC21 | CC22 | CC23 | CC24 | Name: | CC1 | CC2 | CC3 | CC4 | CC5 | CC6 | CC7 | CC8 | CC9 | CC10 | CC11 | CC12 | CC13 | CC14 | CC15 | CC16 | CC17 | CC18 | CC19 | CC20 | CC21 | CC22 | CC23 | CC24
360 | 5.11 | 7.03 | 6.86 | 5.25 | 6.81 | 6.99 | 4.83 | 6.22 | 6.06 | 5.31 | 5.24 | 5.15 | 4.6 | 4.95 | 4.82 | 5.08 | 6.64 | 5.34 | 7.92 | 7.53 | 7.04 | 6.18 | 5.65 | 3.63 | 360 | [24 empty] | 360
370 | 5.49 | 8.13 | 8.23 | 5.53 | 8.32 | 8.35 | 5.11 | 7.26 | 7.11 | 6.32 | 5.61 | 5.63 | 5.11 | 5.31 | 5.18 | 5.47 | 8.08 | 6.2 | 9.54 | 9.01 | 8.57 | 7.31 | 6.34 | 3.96 | 370 | [24 empty] | 370
380 | 5.76 | 9.89 | 11.04 | 5.67 | 11.37 | 10.92 | 5.14 | 9.01 | 8.6 | 8.07 | 5.72 | 5.87 | 5.66 | 5.49 | 5.16 | 5.56 | 10.57 | 7.9 | 12.9 | 11.87 | 11.53 | 9.05 | 7.19 | 4.03 | 380 | [24 empty] | 380
390 | 6.11 | 13.11 | 17.13 | 5.93 | 18.55 | 16.51 | 5.35 | 12.64 | 11.07 | 11.34 | 6.01 | 6.17 | 6.79 | 5.78 | 5.17 | 5.59 | 16.5 | 11.98 | 21.66 | 19.32 | 17.87 | 12.33 | 8.38 | 4.07 | 390 | [24 empty] | 390
400 | 6.42 | 16.24 | 25.83 | 6.1 | 30.01 | 24.4 | 5.55 | 17.65 | 13.2 | 15.31 | 6.25 | 6.47 | 8.96 | 6.02 | 5.19 | 5.69 | 26.61 | 18.21 | 39.95 | 34.2 | 27.36 | 16.11 | 9.39 | 4.14 | 400 | 6.61 | 4.24 | 25.8 | 6.14 | 30.66 | 24.84 | 5.86 | 18.02 | 13.57 | 15.67 | 6.46 | 6.79 | 9.21 | 6.03 | 5.41 | 6.03 | 27.19 | 18.64 | 40.54 | 34.89 | 27.87 | 16.43 | 9.39 | 4.2 | 400 | 6.03 | 15.71 | 24.92 | 5.79 | 29.14 | 23.64 | 5.41 | 17.14 | 12.8 | 14.77 | 6.1 | 6.27 | 8.76 | 5.82 | 5 | 5.59 | 25.72 | 17.56 | 38.68 | 33.38 | 26.77 | 15.64 | 9.02 | 4.01
410 | 6.43 | 17.64 | 31.57 | 6.1 | 38.73 | 29.86 | 5.6 | 23.79 | 13.82 | 17.71 | 6.33 | 6.49 | 13.99 | 6.07 | 5.1 | 5.66 | 33.96 | 23.32 | 64.09 | 49.74 | 33.85 | 17.97 | 9.7 | 4.08 | 410 | 6.67 | 4.19 | 31.07 | 6.18 | 39.19 | 30.09 | 5.86 | 23.71 | 14.07 | 17.93 | 6.52 | 6.76 | 13.91 | 6.1 | 5.29 | 5.92 | 34.31 | 23.64 | 63.87 | 50.13 | 34.14 | 18.19 | 9.69 | 4.19 | 410 | 6.11 | 17.01 | 30.32 | 5.86 | 37.35 | 28.78 | 5.44 | 22.69 | 13.4 | 16.99 | 6.2 | 6.32 | 13.22 | 5.92 | 4.95 | 5.58 | 32.64 | 22.31 | 60.27 | 47.74 | 32.93 | 17.39 | 9.38 | 4
420 | 6.4 | 18.08 | 32.67 | 6.09 | 41.56 | 31.92 | 5.61 | 31.03 | 13.61 | 18.3 | 6.41 | 6.48 | 20.72 | 6.13 | 5.06 | 5.66 | 35.55 | 25.72 | 77.57 | 55.31 | 35.34 | 18.34 | 9.74 | 4.08 | 420 | 6.69 | 4.15 | 32.47 | 6.15 | 41.99 | 32.14 | 5.81 | 31.06 | 13.82 | 18.59 | 6.55 | 6.71 | 20.86 | 6.15 | 5.19 | 5.86 | 35.98 | 25.98 | 77.75 | 55.86 | 35.64 | 18.55 | 9.71 | 4.11 | 420 | 6.05 | 17.42 | 31.57 | 5.85 | 40.31 | 30.92 | 5.42 | 29.58 | 13.21 | 17.65 | 6.25 | 6.28 | 19.62 | 5.97 | 4.87 | 5.53 | 34.44 | 24.74 | 73.89 | 53.66 | 34.6 | 17.85 | 9.41 | 3.97
430 | 6.36 | 18.58 | 32.74 | 6.14 | 42.35 | 33.47 | 5.64 | 36.23 | 13.4 | 17.62 | 6.62 | 6.5 | 25.44 | 6.28 | 5.03 | 5.75 | 34.83 | 28.28 | 81.62 | 56.43 | 35.84 | 18.64 | 9.8 | 4.06 | 430 | 6.65 | 4.09 | 32.72 | 6.13 | 42.82 | 33.66 | 5.8 | 36.4 | 13.52 | 17.86 | 6.73 | 6.68 | 25.67 | 6.27 | 5.13 | 5.88 | 35.21 | 28.49 | 81.88 | 56.97 | 36.02 | 18.79 | 9.72 | 4.06 | 430 | 6.02 | 17.87 | 31.66 | 5.86 | 41.17 | 32.44 | 5.44 | 34.8 | 12.99 | 17.06 | 6.44 | 6.28 | 24.33 | 6.11 | 4.84 | 5.59 | 33.8 | 27.15 | 78.17 | 54.92 | 35.07 | 18.11 | 9.45 | 3.94
440 | 6.3 | 19.37 | 32.55 | 6.15 | 42.44 | 35.24 | 5.66 | 39.38 | 13.19 | 16.11 | 6.96 | 6.52 | 28.78 | 6.54 | 5 | 5.87 | 32.75 | 31.81 | 83.25 | 56.9 | 36.22 | 18.92 | 9.84 | 4.01 | 440 | 6.64 | 4.07 | 32.73 | 6.19 | 42.9 | 35.4 | 5.81 | 39.69 | 13.33 | 16.36 | 7.05 | 6.7 | 29.06 | 6.54 | 5.1 | 5.99 | 33.12 | 31.99 | 83.63 | 57.47 | 36.4 | 19.08 | 9.78 | 4.04 | 440 | 6.03 | 18.64 | 31.49 | 5.92 | 41.33 | 34.19 | 5.5 | 38.06 | 12.82 | 15.67 | 6.78 | 6.34 | 27.64 | 6.39 | 4.85 | 5.75 | 31.87 | 30.48 | 80.01 | 55.45 | 35.45 | 18.39 | 9.51 | 3.92
450 | 6.32 | 20.69 | 31.97 | 6.22 | 41.96 | 37.8 | 5.71 | 40.57 | 12.99 | 14.39 | 7.6 | 6.6 | 30.34 | 7.06 | 5.01 | 6.16 | 30.15 | 35.52 | 84.58 | 57.2 | 36.4 | 19.1 | 9.86 | 3.99 | 450 | 6.63 | 4.06 | 32.31 | 6.25 | 42.4 | 37.97 | 5.85 | 40.93 | 13.12 | 14.55 | 7.68 | 6.77 | 30.75 | 7.08 | 5.1 | 6.29 | 30.44 | 35.76 | 85.01 | 57.77 | 36.61 | 19.23 | 9.77 | 4.02 | 450 | 6.02 | 19.88 | 30.93 | 5.98 | 40.88 | 36.63 | 5.54 | 39.37 | 12.62 | 13.96 | 7.39 | 6.41 | 29.26 | 6.89 | 4.85 | 6.02 | 29.34 | 34.05 | 81.4 | 55.74 | 35.64 | 18.53 | 9.49 | 3.89
460 | 6.33 | 22.88 | 30.89 | 6.29 | 40.62 | 41.7 | 5.77 | 39.89 | 12.65 | 12.53 | 8.84 | 6.68 | 29.1 | 8.12 | 4.96 | 6.73 | 27.32 | 39.17 | 85.56 | 57.12 | 36.24 | 19.02 | 9.75 | 3.97 | 460 | 6.61 | 4.03 | 31.42 | 6.3 | 41.07 | 41.77 | 5.86 | 40.2 | 12.75 | 12.67 | 8.89 | 6.81 | 29.47 | 8.1 | 5.03 | 6.81 | 27.6 | 39.33 | 85.99 | 57.67 | 36.43 | 19.12 | 9.65 | 3.99 | 460 | 6.01 | 21.93 | 29.93 | 6.05 | 39.65 | 40.31 | 5.59 | 38.95 | 12.29 | 12.18 | 8.56 | 6.48 | 28.35 | 7.88 | 4.8 | 6.56 | 26.65 | 37.51 | 82.44 | 55.73 | 35.51 | 18.45 | 9.4 | 3.87
470 | 6.33 | 26.04 | 29.78 | 6.36 | 38.76 | 46.32 | 5.84 | 37.68 | 12.21 | 10.95 | 11 | 6.77 | 26.21 | 10 | 4.9 | 7.73 | 24.79 | 42.51 | 86.27 | 56.94 | 36 | 18.87 | 9.64 | 3.95 | 470 | 6.61 | 4 | 30.48 | 6.37 | 39.14 | 46.46 | 5.92 | 37.97 | 12.29 | 11.05 | 11.05 | 6.88 | 26.54 | 9.96 | 4.94 | 7.82 | 24.99 | 42.77 | 86.73 | 57.5 | 36.17 | 18.96 | 9.53 | 3.97 | 470 | 6 | 24.98 | 28.82 | 6.11 | 37.78 | 44.85 | 5.66 | 36.85 | 11.85 | 10.6 | 10.67 | 6.57 | 25.61 | 9.72 | 4.73 | 7.56 | 24.09 | 40.76 | 83.13 | 55.57 | 35.26 | 18.31 | 9.28 | 3.85
480 | 6.37 | 30.11 | 28.64 | 6.49 | 36.22 | 50.92 | 6 | 33.26 | 11.7 | 9.55 | 14.95 | 6.94 | 22.23 | 13.38 | 4.84 | 9.89 | 21.85 | 44.15 | 86.91 | 56.83 | 35.81 | 18.76 | 9.54 | 3.94 | 480 | 6.64 | 3.98 | 29.58 | 6.48 | 36.68 | 50.89 | 6.07 | 33.78 | 11.77 | 9.68 | 14.8 | 7.06 | 22.69 | 13.17 | 4.88 | 9.85 | 22.14 | 44.42 | 87.36 | 57.32 | 35.92 | 18.84 | 9.44 | 3.95 | 480 | 6.04 | 28.72 | 27.77 | 6.22 | 35.42 | 49.21 | 5.8 | 32.77 | 11.37 | 9.28 | 14.25 | 6.74 | 21.87 | 12.79 | 4.68 | 9.53 | 21.38 | 42.4 | 83.84 | 55.46 | 35.07 | 18.2 | 9.2 | 3.84
490 | 6.43 | 32.9 | 27.44 | 6.69 | 33.3 | 53.86 | 6.25 | 27.71 | 11.2 | 8.46 | 19.87 | 7.21 | 18.12 | 17.44 | 4.79 | 13.53 | 19.24 | 43.68 | 87.35 | 56.76 | 35.65 | 18.68 | 9.46 | 3.92 | 490 | 6.7 | 3.95 | 28.56 | 6.7 | 33.56 | 54.19 | 6.35 | 27.8 | 11.27 | 8.49 | 20.14 | 7.36 | 18.25 | 17.6 | 4.84 | 13.76 | 19.31 | 43.97 | 87.88 | 57.25 | 35.8 | 18.77 | 9.36 | 3.92 | 490 | 6.11 | 31.65 | 26.56 | 6.46 | 32.48 | 52.32 | 6.07 | 27.11 | 10.87 | 8.14 | 19.26 | 7.05 | 17.65 | 16.93 | 4.65 | 13.24 | 18.7 | 42.03 | 84.4 | 55.38 | 34.94 | 18.14 | 9.13 | 3.81
500 | 6.67 | 34.73 | 25.9 | 7.42 | 30.07 | 55.56 | 6.95 | 21.84 | 10.68 | 7.51 | 27.54 | 8.15 | 13.84 | 23.15 | 4.85 | 21.13 | 17.09 | 41.28 | 87.82 | 56.82 | 35.66 | 18.69 | 9.42 | 3.91 | 500 | 6.93 | 3.95 | 27.31 | 7.44 | 30.38 | 55.88 | 7.05 | 22.04 | 10.77 | 7.56 | 27.71 | 8.31 | 14.06 | 23.22 | 4.9 | 21.3 | 17.2 | 41.6 | 88.35 | 57.3 | 35.81 | 18.8 | 9.34 | 3.92 | 500 | 6.34 | 33.39 | 25.13 | 7.24 | 29.44 | 54.05 | 6.81 | 21.6 | 10.41 | 7.25 | 26.57 | 8.08 | 13.68 | 22.35 | 4.72 | 20.57 | 16.64 | 39.82 | 84.87 | 55.47 | 34.97 | 18.17 | 9.12 | 3.8
510 | 7.3 | 34.71 | 24.41 | 10.24 | 27.57 | 55.89 | 9.23 | 17.2 | 10.17 | 6.91 | 37.94 | 11.85 | 10.51 | 29.36 | 5.09 | 33.78 | 15.34 | 37.3 | 88.11 | 56.95 | 35.73 | 18.77 | 9.43 | 3.9 | 510 | 7.54 | 3.94 | 26 | 10.24 | 27.86 | 56.18 | 9.35 | 17.44 | 10.25 | 6.95 | 38.03 | 12.05 | 10.73 | 29.33 | 5.13 | 33.82 | 15.46 | 37.62 | 88.73 | 57.42 | 35.89 | 18.87 | 9.34 | 3.91 | 510 | 6.91 | 33.36 | 23.68 | 9.88 | 26.95 | 54.38 | 8.99 | 17.06 | 9.91 | 6.65 | 36.33 | 11.7 | 10.43 | 28.19 | 4.92 | 32.44 | 14.92 | 36.07 | 85.28 | 55.59 | 35.04 | 18.23 | 9.12 | 3.79
520 | 7.96 | 31.36 | 22.89 | 15.3 | 25.28 | 55.14 | 13.16 | 13.68 | 9.7 | 6.48 | 47.39 | 20.93 | 8.04 | 33.1 | 5.27 | 47.24 | 13.05 | 32.27 | 88.3 | 57.1 | 35.84 | 18.83 | 9.44 | 3.88 | 520 | 8.18 | 3.92 | 24.53 | 15.3 | 25.53 | 55.41 | 13.3 | 13.8 | 9.76 | 6.51 | 47.65 | 20.92 | 8.17 | 33.24 | 5.31 | 47.42 | 13.17 | 32.55 | 88.9 | 57.49 | 35.93 | 18.91 | 9.33 | 3.89 | 520 | 7.54 | 30.37 | 22.19 | 14.6 | 24.69 | 53.7 | 12.7 | 13.5 | 9.44 | 6.21 | 45.61 | 20.14 | 7.93 | 31.98 | 5.09 | 45.53 | 12.75 | 31.23 | 85.52 | 55.7 | 35.1 | 18.29 | 9.09 | 3.77
530 | 8.21 | 27.13 | 21.11 | 18.47 | 23.08 | 53.45 | 16.11 | 11.28 | 9.44 | 6.05 | 52.61 | 31.8 | 6.19 | 33.73 | 5.27 | 56.22 | 10.97 | 26.92 | 88.5 | 57.12 | 35.83 | 18.84 | 9.39 | 3.85 | 530 | 8.43 | 3.9 | 22.85 | 18.59 | 23.32 | 53.78 | 16.3 | 11.39 | 9.52 | 6.08 | 52.99 | 31.96 | 6.29 | 33.94 | 5.32 | 56.59 | 11.06 | 27.21 | 89.1 | 57.6 | 36 | 18.96 | 9.31 | 3.86 | 530 | 7.83 | 26.39 | 20.52 | 17.77 | 22.57 | 52.13 | 15.63 | 11.1 | 9.19 | 5.81 | 51.02 | 30.58 | 6.08 | 32.75 | 5.11 | 54.53 | 10.74 | 26.14 | 85.79 | 55.79 | 35.15 | 18.33 | 9.08 | 3.75
540 | 8.31 | 25.62 | 19.57 | 18.38 | 21.83 | 50.98 | 17.7 | 10.05 | 9.39 | 5.77 | 53.98 | 39.39 | 5.12 | 32.26 | 5.2 | 60.89 | 10.3 | 22.04 | 88.68 | 57.08 | 35.78 | 18.83 | 9.35 | 3.84 | 540 | 8.53 | 3.88 | 21.31 | 18.5 | 22.04 | 51.31 | 17.88 | 10.12 | 9.48 | 5.79 | 54.54 | 39.79 | 5.17 | 32.49 | 5.26 | 61.46 | 10.34 | 22.25 | 89.37 | 57.62 | 35.97 | 18.97 | 9.28 | 3.85 | 540 | 7.96 | 24.83 | 19.04 | 17.84 | 21.32 | 49.77 | 17.27 | 9.84 | 9.16 | 5.54 | 52.72 | 38.11 | 4.99 | 31.45 | 5.06 | 59.38 | 10.04 | 21.44 | 86.06 | 55.82 | 35.15 | 18.35 | 9.06 | 3.74
550 | 8.59 | 26.25 | 18.43 | 16.76 | 21.29 | 47.71 | 19.8 | 9.32 | 9.47 | 5.72 | 53.58 | 43.49 | 4.55 | 29.64 | 5.14 | 63.8 | 10.64 | 17.68 | 88.96 | 57.09 | 35.77 | 18.85 | 9.34 | 3.84 | 550 | 8.77 | 3.87 | 20.11 | 16.81 | 21.5 | 48.01 | 19.95 | 9.36 | 9.55 | 5.73 | 54.05 | 43.9 | 4.59 | 29.82 | 5.19 | 64.4 | 10.73 | 17.83 | 89.63 | 57.63 | 35.95 | 18.97 | 9.26 | 3.85 | 550 | 8.21 | 25.37 | 17.9 | 16.31 | 20.76 | 46.6 | 19.33 | 9.09 | 9.25 | 5.48 | 52.37 | 42.18 | 4.41 | 28.92 | 5.01 | 62.32 | 10.32 | 17.22 | 86.35 | 55.84 | 35.13 | 18.36 | 9.04 | 3.74
560 | 9.27 | 26.33 | 17.22 | 14.65 | 20.57 | 43.66 | 24.56 | 8.66 | 9.69 | 5.76 | 51.77 | 46.53 | 4.24 | 26.32 | 5.2 | 65.96 | 10.77 | 13.9 | 88.99 | 57.17 | 35.8 | 18.91 | 9.33 | 3.83 | 560 | 9.4 | 3.86 | 18.93 | 14.74 | 20.77 | 44.04 | 24.56 | 8.72 | 9.8 | 5.79 | 52.28 | 46.82 | 4.28 | 26.57 | 5.27 | 66.49 | 10.85 | 14.13 | 89.7 | 57.63 | 35.96 | 19 | 9.24 | 3.84 | 560 | 8.85 | 25.55 | 16.75 | 14.34 | 20.09 | 42.76 | 23.83 | 8.47 | 9.53 | 5.52 | 50.73 | 45.11 | 4.11 | 25.8 | 5.09 | 64.46 | 10.5 | 13.65 | 86.52 | 55.88 | 35.16 | 18.4 | 9.03 | 3.74
570 | 10.55 | 27.37 | 15.8 | 13.04 | 19.7 | 39.28 | 32.9 | 8.28 | 10.73 | 5.71 | 49.21 | 50.12 | 4.11 | 22.9 | 5.72 | 67.63 | 11.34 | 11.37 | 89.02 | 57.27 | 35.9 | 19.02 | 9.35 | 3.83 | 570 | 10.63 | 3.85 | 17.59 | 13.08 | 19.86 | 39.56 | 32.96 | 8.31 | 10.84 | 5.73 | 49.65 | 50.42 | 4.14 | 23.1 | 5.81 | 68.18 | 11.39 | 11.54 | 89.68 | 57.7 | 36.01 | 19.1 | 9.25 | 3.82 | 570 | 10.02 | 26.51 | 15.39 | 12.73 | 19.23 | 38.43 | 31.85 | 8.06 | 10.6 | 5.48 | 48.22 | 48.58 | 3.97 | 22.45 | 5.6 | 66.11 | 11.06 | 11.13 | 86.58 | 55.95 | 35.22 | 18.48 | 9.03 | 3.72
580 | 12.23 | 32.94 | 14.54 | 12.08 | 19.41 | 34.35 | 43.53 | 8.23 | 15.99 | 5.58 | 45.85 | 54.23 | 4.06 | 19.25 | 7.23 | 68.93 | 14.52 | 9.84 | 89.02 | 57.26 | 35.89 | 19.05 | 9.32 | 3.81 | 580 | 12.31 | 3.84 | 16.45 | 12.11 | 19.57 | 34.62 | 43.84 | 8.25 | 15.98 | 5.6 | 46.28 | 54.66 | 4.09 | 19.41 | 7.32 | 69.57 | 14.54 | 9.98 | 89.75 | 57.75 | 36.08 | 19.17 | 9.24 | 3.82 | 580 | 11.64 | 31.65 | 14.17 | 11.77 | 18.94 | 33.68 | 42.25 | 8 | 15.51 | 5.37 | 45.01 | 52.65 | 3.92 | 18.92 | 7.08 | 67.49 | 14.03 | 9.6 | 86.67 | 56.01 | 35.28 | 18.56 | 9.03 | 3.71
590 | 13.75 | 41.41 | 13.82 | 11.47 | 19.99 | 29.2 | 53.14 | 8.35 | 27.57 | 5.62 | 41.66 | 57.85 | 4.06 | 15.62 | 10.96 | 69.88 | 20.89 | 8.89 | 89.01 | 57.18 | 35.81 | 19.03 | 9.28 | 3.8 | 590 | 13.84 | 3.84 | 15.8 | 11.51 | 20.12 | 29.44 | 53.69 | 8.39 | 27.59 | 5.62 | 42.1 | 58.36 | 4.09 | 15.74 | 10.94 | 70.58 | 20.92 | 9.02 | 89.79 | 57.69 | 36.02 | 19.16 | 9.22 | 3.81 | 590 | 13.13 | 39.83 | 13.45 | 11.15 | 19.46 | 28.67 | 51.85 | 8.12 | 26.6 | 5.41 | 40.99 | 56.25 | 3.91 | 15.38 | 10.66 | 68.51 | 20.14 | 8.64 | 86.76 | 55.96 | 35.23 | 18.56 | 9 | 3.7
600 | 14.55 | 48.5 | 13.61 | 10.91 | 21.29 | 24.61 | 58.63 | 8.48 | 40.99 | 6.18 | 37.33 | 60.18 | 4.07 | 12.66 | 18.73 | 70.43 | 29.6 | 8.16 | 88.99 | 56.97 | 35.61 | 18.92 | 9.22 | 3.8 | 600 | 14.63 | 3.83 | 15.63 | 10.94 | 21.46 | 24.79 | 59.2 | 8.52 | 41.3 | 6.18 | 37.72 | 60.76 | 4.1 | 12.73 | 18.71 | 71.1 | 29.69 | 8.28 | 89.75 | 57.49 | 35.82 | 19.06 | 9.15 | 3.81 | 600 | 13.96 | 46.87 | 13.24 | 10.63 | 20.72 | 24.2 | 57.4 | 8.25 | 39.73 | 5.95 | 36.79 | 58.62 | 3.93 | 12.49 | 18.13 | 69.07 | 28.58 | 7.95 | 86.79 | 55.79 | 35.05 | 18.48 | 8.95 | 3.71
610 | 14.79 | 52.91 | 13.7 | 10.45 | 22.72 | 21.47 | 60.23 | 8.52 | 51.03 | 7.64 | 33.94 | 61.39 | 4.12 | 10.83 | 31.31 | 70.71 | 40.72 | 7.64 | 88.95 | 56.63 | 35.3 | 18.74 | 9.13 | 3.8 | 610 | 14.86 | 3.82 | 15.82 | 10.49 | 22.9 | 21.71 | 60.7 | 8.56 | 51.22 | 7.65 | 34.4 | 62 | 4.13 | 10.93 | 31.4 | 71.44 | 40.86 | 7.76 | 89.69 | 57.19 | 35.54 | 18.89 | 9.06 | 3.8 | 610 | 14.23 | 51.23 | 13.32 | 10.19 | 22.1 | 21.19 | 59.09 | 8.3 | 49.42 | 7.32 | 33.57 | 59.87 | 3.96 | 10.72 | 30.1 | 69.44 | 39.2 | 7.45 | 86.78 | 55.54 | 34.8 | 18.32 | 8.85 | 3.7
620 | 15.23 | 55.42 | 13.91 | 10.4 | 23.84 | 20.08 | 60.29 | 8.57 | 55.69 | 10.11 | 32.3 | 62.02 | 4.16 | 10.09 | 45.62 | 71.3 | 53.41 | 7.43 | 89.01 | 56.34 | 35.03 | 18.58 | 9.05 | 3.79 | 620 | 15.26 | 3.82 | 16.3 | 10.38 | 23.98 | 20.19 | 60.8 | 8.57 | 56.22 | 10.04 | 32.62 | 62.67 | 4.18 | 10.11 | 45.74 | 72.01 | 53.49 | 7.51 | 89.77 | 56.9 | 35.23 | 18.7 | 8.97 | 3.79 | 620 | 14.6 | 53.84 | 13.5 | 10.07 | 23.2 | 19.66 | 59.26 | 8.32 | 54.49 | 9.59 | 31.78 | 60.56 | 3.99 | 9.86 | 43.88 | 70.02 | 51.4 | 7.2 | 86.89 | 55.25 | 34.51 | 18.14 | 8.76 | 3.68
630 | 16.13 | 57.28 | 14.03 | 10.61 | 25.28 | 19.39 | 60.19 | 8.83 | 57.65 | 12.81 | 31.46 | 62.48 | 4.21 | 9.71 | 56.02 | 72.03 | 64.5 | 7.31 | 89.18 | 56.01 | 34.71 | 18.39 | 8.95 | 3.77 | 630 | 16.17 | 3.8 | 16.91 | 10.59 | 25.4 | 19.47 | 60.72 | 8.83 | 58.25 | 12.78 | 31.76 | 63.12 | 4.24 | 9.73 | 56.42 | 72.78 | 64.82 | 7.41 | 89.92 | 56.58 | 34.91 | 18.5 | 8.88 | 3.78 | 630 | 15.44 | 55.76 | 13.63 | 10.26 | 24.58 | 18.93 | 59.19 | 8.58 | 56.62 | 12.25 | 30.92 | 61 | 4.06 | 9.48 | 54.34 | 70.77 | 62.49 | 7.11 | 87.06 | 54.98 | 34.2 | 17.97 | 8.67 | 3.68
640 | 17.35 | 59.14 | 14.15 | 10.95 | 27.67 | 18.95 | 60.07 | 9.62 | 58.42 | 15.43 | 30.94 | 62.86 | 4.33 | 9.48 | 61.46 | 72.9 | 72.29 | 7.3 | 89.38 | 55.74 | 34.41 | 18.2 | 8.86 | 3.77 | 640 | 17.42 | 3.8 | 17.53 | 10.95 | 27.78 | 19.04 | 60.54 | 9.61 | 59.06 | 15.44 | 31.23 | 63.5 | 4.35 | 9.51 | 62.09 | 73.63 | 72.87 | 7.39 | 90.09 | 56.27 | 34.6 | 18.33 | 8.8 | 3.77 | 640 | 16.64 | 57.59 | 13.77 | 10.63 | 26.88 | 18.52 | 59.07 | 9.34 | 57.48 | 14.86 | 30.4 | 61.42 | 4.19 | 9.29 | 60 | 71.63 | 70.49 | 7.11 | 87.26 | 54.71 | 33.92 | 17.82 | 8.62 | 3.7
650 | 18.53 | 61.14 | 14.45 | 11.24 | 30.52 | 18.7 | 59.96 | 10.98 | 58.64 | 18.11 | 30.64 | 63.19 | 4.49 | 9.39 | 64.03 | 73.66 | 77.11 | 7.42 | 89.45 | 55.41 | 34.09 | 18.02 | 8.78 | 3.77 | 650 | 18.62 | 3.79 | 18.19 | 11.24 | 30.68 | 18.76 | 60.48 | 10.95 | 59.39 | 18.13 | 30.89 | 63.94 | 4.51 | 9.39 | 64.8 | 74.43 | 77.86 | 7.49 | 90.31 | 56.01 | 34.32 | 18.16 | 8.71 | 3.77 | 650 | 17.79 | 59.66 | 14.06 | 10.92 | 29.69 | 18.28 | 59.04 | 10.63 | 57.85 | 17.51 | 30.11 | 61.85 | 4.34 | 9.19 | 62.79 | 72.44 | 75.5 | 7.2 | 87.49 | 54.47 | 33.65 | 17.66 | 8.54 | 3.69
660 | 19.29 | 63.51 | 14.84 | 11.4 | 32.88 | 18.97 | 60.57 | 12.96 | 58.8 | 21.13 | 30.98 | 63.6 | 4.66 | 9.54 | 65.35 | 74.28 | 79.68 | 7.6 | 89.39 | 55.16 | 33.8 | 17.84 | 8.68 | 3.75 | 660 | 19.45 | 3.79 | 18.84 | 11.43 | 33.16 | 19.02 | 61 | 12.93 | 59.52 | 21.2 | 31.22 | 64.3 | 4.71 | 9.55 | 66.14 | 75.07 | 80.56 | 7.7 | 90.26 | 55.72 | 34.03 | 17.99 | 8.64 | 3.77 | 660 | 18.58 | 61.96 | 14.47 | 11.1 | 32.11 | 18.5 | 59.57 | 12.54 | 58.02 | 20.48 | 30.38 | 62.25 | 4.52 | 9.33 | 64.17 | 73.1 | 78.27 | 7.39 | 87.54 | 54.22 | 33.39 | 17.51 | 8.47 | 3.69
670 | 19.25 | 66.06 | 14.93 | 11.35 | 33.69 | 19.91 | 61.9 | 15.69 | 58.81 | 24.82 | 32.18 | 63.96 | 4.73 | 10.03 | 65.99 | 74.74 | 80.8 | 7.69 | 89.14 | 54.88 | 33.5 | 17.66 | 8.59 | 3.73 | 670 | 19.48 | 3.79 | 19.08 | 11.4 | 34.1 | 19.93 | 62.33 | 15.64 | 59.5 | 24.89 | 32.39 | 64.62 | 4.82 | 10.04 | 66.8 | 75.51 | 81.68 | 7.82 | 89.89 | 55.39 | 33.72 | 17.8 | 8.56 | 3.77 | 670 | 18.61 | 64.41 | 14.58 | 11.07 | 33.04 | 19.35 | 60.78 | 15.16 | 58 | 24.05 | 31.45 | 62.52 | 4.61 | 9.77 | 64.82 | 73.49 | 79.43 | 7.5 | 87.19 | 53.88 | 33.08 | 17.32 | 8.38 | 3.67
680 | 18.42 | 69.07 | 14.57 | 11.06 | 32.92 | 21.22 | 63.8 | 19.5 | 58.84 | 29.66 | 33.92 | 64.45 | 4.68 | 10.76 | 66.59 | 75.46 | 81.54 | 7.62 | 89.18 | 54.64 | 33.2 | 17.47 | 8.51 | 3.73 | 680 | 18.69 | 3.78 | 18.81 | 11.11 | 33.31 | 21.19 | 64.06 | 19.36 | 59.44 | 29.6 | 34.05 | 64.97 | 4.76 | 10.74 | 67.3 | 76.08 | 82.37 | 7.74 | 89.82 | 55.05 | 33.4 | 17.61 | 8.48 | 3.75 | 680 | 17.95 | 67.25 | 14.29 | 10.83 | 32.47 | 20.57 | 62.5 | 18.74 | 58.02 | 28.61 | 33.07 | 62.94 | 4.58 | 10.44 | 65.41 | 74.14 | 80.2 | 7.45 | 87.24 | 53.63 | 32.82 | 17.17 | 8.32 | 3.67
690 | 17.59 | 71.55 | 13.89 | 10.7 | 31.67 | 22.06 | 64.87 | 23.63 | 58.47 | 34.57 | 35.06 | 64.47 | 4.63 | 11.27 | 66.82 | 75.99 | 82.19 | 7.39 | 89.36 | 54.01 | 32.64 | 17.15 | 8.38 | 3.71 | 690 | 18 | 3.78 | 18.43 | 10.85 | 32.31 | 22.42 | 65.81 | 24.03 | 59.59 | 35.26 | 35.74 | 65.56 | 4.72 | 11.46 | 68.06 | 77.1 | 83.41 | 7.59 | 90.32 | 54.89 | 33.19 | 17.49 | 8.43 | 3.76 | 690 | 17.26 | 70.27 | 13.78 | 10.59 | 31.5 | 21.76 | 64.15 | 23.3 | 58.09 | 34.14 | 34.68 | 63.42 | 4.57 | 11.15 | 66.05 | 75.01 | 81.12 | 7.32 | 87.56 | 53.42 | 32.57 | 17.04 | 8.27 | 3.68
700 | 17.54 | 74.08 | 13.31 | 10.67 | 31.49 | 22.86 | 65.96 | 28.33 | 58.54 | 39.98 | 36.2 | 64.87 | 4.73 | 11.79 | 67.34 | 76.67 | 83 | 7.19 | 89.61 | 53.78 | 32.43 | 17.03 | 8.35 | 3.73 | 700 | 18.02 | 3.79 | 17.9 | 10.82 | 32.19 | 23.27 | 67.05 | 29.69 | 59.65 | 41.75 | 36.99 | 66.03 | 4.85 | 12.01 | 68.66 | 77.89 | 84.33 | 7.34 | 90.54 | 54.6 | 32.93 | 17.34 | 8.36 | 3.76 | 700 | 17.23 | 73.17 | 13.13 | 10.55 | 31.34 | 22.62 | 65.43 | 28.76 | 58.21 | 40.39 | 35.94 | 63.93 | 4.68 | 11.69 | 66.7 | 75.84 | 82.06 | 7.09 | 87.88 | 53.2 | 32.35 | 16.9 | 8.21 | 3.68

## CALCULATION
Inst, no: | Conditions: | Name: | X | Y | Z | x | y | Wcie | Tcie | Wgg | Tgg | L* | a* | b* | dL* | da* | db* | C* | dC* | hab (rad) | hab (deg) | dhab | dH* | dE*
435 | LAV I Calib | cc1 | 11.11 | 10.05 | 6.84 | 0.3967 | 0.359 | -103.9 | -56.37 | -244 | -63.18 | 37.94 | 12.17 | 13.11 |  |  |  | 17.888 |  | 0.82256 | 47.1295
596 | LAV I Calib | CC1 | 11.24 | 10.22 | 7.15 | 0.3928 | 0.3572 | -97.47 | -54.1 | -230.5 | -60.86 | 38.23 | 11.86 | 12.41 | 0.29 | -0.31 | -0.7 | 17.1659 | -0.72213 | 0.80806 | 46.2982 | -0.014509 | -0.25423 | 0.81866
Inst, no: | Conditions: | Name: | X | Y | Z | x | y | Wcie | Tcie | Wgg | Tgg | L* | a* | b* | dL* | da* | db* | C* | dC* | hab (rad) | hab (deg) | dhab | dH* | dE*
435 | LAV I Calib | cc1 | 11.11 | 10.05 | 6.84 | 0.3967 | 0.359 | -103.9 | -56.37 | -244 | -63.18 | 37.94 | 12.17 | 13.11 |  |  |  | 17.888 |  | 0.82256 | 47.1295
118 | %R LAV SCI Calib | CC1 | 10.62 | 9.6 | 6.5 | 0.3976 | 0.3592 | -105.3 | -57.09 |  |  | 37.11 | 12.13 | 13.03 | -0.83 | -0.04 | -0.08 | 17.8022 | -0.085827 | 0.82115 | 47.0487 | -0.0014106 | -0.025172 | 0.83481
Inst, no: | Conditions: | Name: | X | Y | Z | x | y | Wcie | Tcie | Wgg | Tgg | L* | a* | b* | dL* | da* | db* | C* | dC* | hab (rad) | hab (deg) | dhab | dH* | dE*
596 | LAV I Calib | CC1 | 11.24 | 10.22 | 7.15 | 0.3928 | 0.3572 | -97.47 | -54.1 | -230.5 | -60.86 | 38.23 | 11.86 | 12.41 |  |  |  | 17.1659 |  | 0.80806 | 46.2982
118 | %R LAV SCI Calib | CC1 | 10.62 | 9.6 | 6.5 | 0.3976 | 0.3592 | -105.3 | -57.09 |  |  | 37.11 | 12.13 | 13.03 | -1.12 | 0.27 | 0.62 | 17.8022 | 0.6363 | 0.82115 | 47.0487 | 0.013098 | 0.22897 | 1.3083
1) SAMPLE CALCULATION OF DIFFERENCES
2) THE FIRST ROW IS THE DATA FOR STANDARD, THE SECOND ROW IS THE DATA FOR BATCH
3) THE DIFFERENCE IS ALWAYS CALCULATED AS THE DIFFERENCE BETWEEN THE SAMPLE AND THE STANDARD
THE PROCEDURE FOR ONE SAMPLE IS SHOWN FOR ILLUSTRATION:
(A) COMPARISON OF THE SF450 SPECTROPHOTOMETER (STANDARD) AND THE SF600 SPECTROPHOTOMETER (BATCH)
(B) COMPARISON OF SPECTROPHOTOMETER SF450 (STANDARD) AND SPECTROPHOTOMETER SF300 (BATCH)
(C) COMPARISON OF THE SF600 (STANDARD) SPECTROPHOTOMETER AND THE SF300 (BATCH) SPECTROPHOTOMETER